HIMMS and RSNA

IHE

Integrating the Healthcare Enterprise

Test Plan for Image Creator

2000-06-21

Marco Eichelberg1,  Klaus Kleber2, Jörg Riesmeier1

1  Kuratorium OFFIS e.V., Escherweg 2, 26121 Oldenburg, Germany

2  Institute for Microtherapy, Universitätsstraße 142, 44799 Bochum, Germany

Contents

1	Introduction	3

2	Image Creator Application Specification	5

2.1	Association Negotiation	5

2.1.1	Association Initiation Policy	5

2.1.2	Application Identification	6

2.1.3	Proposed Abstract Syntax Table	6

2.1.4	Proposed Transfer Syntax Table	8

2.1.5	Other Association Negotiation Items	9

2.2	Data Structure transmitted with the Storage Service SOP Classes	10

2.2.1	Modality LUT Transformation	10

2.2.2	Mask Transformation	10

2.2.3	VOI LUT Transformation	11

2.2.4	Presentation LUT Transformation	11

2.2.5	Overlay	11

2.2.6	Curve	12

2.2.7	Image Pixel related Attributes	13

2.3	Data Structure transmitted with the Grayscale Softcopy Presentation State Storage SOP Class	15

2.3.1	Modality LUT Transformation	15

2.3.2	Mask Transformation	15

2.3.3	VOI LUT Transformation	16

2.3.4	Presentation LUT Transformation	16

2.3.5	Shutter Transformation	16

2.3.6	Image Annotation	17

2.3.7	Spatial Transformation	18

2.3.8	Overlay	18

2.3.9	Curve	19

2.3.10	Displayed Area	20

2.3.11	Other Attributes	20

2.4	Behavior of other SOP Classes	20

3	Test Procedure for Image Displays	22

# Introduction

This document is intended for vendors participating in the IHE Year 2 demonstration which will supply one or more Image Creator Actors as defined in section 3.1 of the IHE Technical Framework (Year 2).

Image Creators are systems that create additional images and/or Grayscale Softcopy Presentation States and transmit the data to an Image Archive. They also make requests for the storage commitment of the Image Manager for the images and/or Presentation States previously transmitted and generate Performed Procedure Steps. The Image Creator must be grouped with an Image Display. They are required to implement the following IHE transactions:

- “Query Image” transaction defined in section 6.14 of the Technical Framework.

- “Retrieve Image” transaction defined in section 6.16 of the Technical Framework.

- “Creator Images Stored” transaction defined in section 6.18 of the Technical Framework.

- “Creator Presentation State Stored” transaction defined in section 6.19 of the Technical Framework.

- “Creator Storage Commitment” transaction defined in section 6.22 of the Technical Framework.

They can optionally support the following IHE transactions:

- “Query Presentation State” transaction defined in section 6.15 of the Technical Framework.

- “Retrieve Presentation State” transaction defined in section 6.17 of the Technical Framework.

- “Creator Procedure Step in Progress” transaction defined in section 6.20 of the Technical Framework.

- “Creator Procedure Step Completed” transaction defined in section 6.21 of the Technical Framework.

This document contains an Application Specification Form which allows the description of the association and negotiation behavior and the data structure transmitted of an Image Creator in terms of DICOM services, messages and objects.  It is roughly structured like a DICOM Conformance Statement but goes into more detail than a typical conformance statement.

This form should be filled in by all vendors supplying an Image Creator.  The filled-in forms will be made available to vendors supplying actors communicating with Image Creators.  In this way the identification of certain DICOM protocol and data structure related problems will be possible at an early point of time.

If a vendor supplies more than one Image Creator, or if the Image Creator can be operated in different modes which exhibit a different behavior on the DICOM network protocol level then multiple copies of the form should be filled in.  Please send the completed forms to Steve Moore at the Mallinckrodt Institute of Radiology:

email:	moores@mir.wustl.edu

fax: 314-362-6971

phone: 314-362-6965

If you send the forms as Word attachments, please run a virus scanner on the document before you send it.  The forms should be completed and returned to Steve Moore by INSERT_DEADLINE_HERE.

If a DICOM conformance statement for the Image Creator is available, it would be useful to provide it together with (but not instead of) the filled-in specification form.

IHE Year 2 Image Creator Actors

Application Specification Form

# Image Creator Application Specification

Application vendor

Name and software Version of the application

Person filling in this form, place, date

## Association Negotiation

This section describes the association negotiation behavior of the application.

### Association Initiation Policy

Please describe the real world activity leading to an association initiation by the application.

Association initiation policy

### Application Identification

Implementation Class UID

Implementation Version Name

Calling Application Entity Title

Max PDU Size

General Comments on the Application:

Comments

### Proposed Abstract Syntax Table

Please check the SOP Classes / Meta SOP Classes that are proposed by the application during the association negotiation (A-ASSOCIATE-RQ):

| Yes | SOP Class Name | SOP Class UID |
| --- | --- | --- |
|  | Computed Radiography Image Storage | 1.2.840.10008.5.1.4.1.1.1 |
|  | CT Image Storage | 1.2.840.10008.5.1.4.1.1.2 |
|  | MR Image Storage | 1.2.840.10008.5.1.4.1.1.4 |
|  | Nuclear Medicine Image Storage | 1.2.840.10008.5.1.4.1.1.20 |
|  | Positron Emission Tomography Image Storage | 1.2.840.10008.5.1.4.1.1.128 |
|  | RT Image Storage | 1.2.840.10008.5.1.4.1.1.481.1 |
|  | Secondary Capture Image Storage | 1.2.840.10008.5.1.4.1.1.7 |
|  | Ultrasound Image Storage | 1.2.840.10008.5.1.4.1.1.6.1 |
|  | Ultrasound Image Storage | 1.2.840.10008.5.1.4.1.1.6.1 |
|  | Ultrasound Multi-frame Image Storage | 1.2.840.10008.5.1.4.1.1.3.1 |
|  | X-Ray Angiographic Image Storage | 1.2.840.10008.5.1.4.1.1.12.1 |
|  | X-Ray Radiofluoroscopic Image Storage | 1.2.840.10008.5.1.4.1.1.12.2 |
|  | Digital X-Ray Image Storage – For Presentation | 1.2.840.10008.5.1.4.1.1.1.1 |
|  | Digital X-Ray Image Storage – For Processing | 1.2.840.10008.5.1.4.1.1.1.1.1 |
|  | Digital Mammography Image Storage – For Presentation | 1.2.840.10008.5.1.4.1.1.1.2 |
|  | Digital Mammography Image Storage – For Processing | 1.2.840.10008.5.1.4.1.1.1.2.1 |
|  | Digital Intra-oral X-Ray Image Storage – For Presentation | 1.2.840.10008.5.1.4.1.1.1.3 |
|  | Digital Intra-oral X-Ray Image Storage – For Processing | 1.2.840.10008.5.1.4.1.1.1.3.1 |
|  | VL Endoscopic Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.1 |
|  | VL Microscopic Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.2 |
|  | VL Slide-Coordinates Microscopic Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.3 |
|  | VL Photographic Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.4 |
|  | Grayscale Softcopy Presentation State Storage | 1.2.840.10008.5.1.4.1.1.11.1 |
|  | Verification | 1.2.840.10008.1.1 |
|  | The following Storage SOP Classes are not required by the IHE Technical Framework: |  |
|  | Hardcopy Color Image Storage | 1.2.840.10008.5.1.1.30 |
|  | Hardcopy Grayscale Image Storage | 1.2.840.10008.5.1.1.29 |
|  | PET Curve Storage | 1.2.840.10008.5.1.4.1.1.129 |
|  | RETIRED Nuclear Medicine Image Storage | 1.2.840.10008.5.1.4.1.1.5 |
|  | RETIRED Ultrasound Image Storage | 1.2.840.10008.5.1.4.1.1.6 |
|  | RETIRED Ultrasound Multiframe Image Storage | 1.2.840.10008.5.1.4.1.1.3 |
|  | RETIRED X-ray Angiographic Bi-Plane Image Storage | 1.2.840.10008.5.1.4.1.1.12.3 |
|  | RT Beams Treatment Record Storage | 1.2.840.10008.5.1.4.1.1.481.4 |
|  | RT Brachy Treatment Record Storage | 1.2.840.10008.5.1.4.1.1.481.6 |
|  | RT Dose Storage | 1.2.840.10008.5.1.4.1.1.481.2 |
|  | RT Plan Storage | 1.2.840.10008.5.1.4.1.1.481.5 |
|  | RT Structure Set Storage | 1.2.840.10008.5.1.4.1.1.481.3 |
|  | RT Treatment Summary Record Storage | 1.2.840.10008.5.1.4.1.1.481.7 |
|  | Standalone Curve Storage | 1.2.840.10008.5.1.4.1.1.9 |
|  | Standalone Modality LUT Storage | 1.2.840.10008.5.1.4.1.1.10 |
|  | Standalone Overlay Storage | 1.2.840.10008.5.1.4.1.1.8 |
|  | Standalone VOI LUT Storage | 1.2.840.10008.5.1.4.1.1.11 |
|  | Stored Print Storage | 1.2.840.10008.5.1.1.27 |
|  | Other (please specify): |  |

General Comments on the Proposed Abstract Syntax Table:

Comments

### Proposed Transfer Syntax Table

Please check the Transfer Syntaxes that are proposed by the application during the association negotiation (A-ASSOCIATE-RQ) for each Abstract Syntax:

| Yes | Transfer Syntax Name | Transfer Syntax UID |
| --- | --- | --- |
|  | Implicit VR Little Endian | 1.2.840.10008.1.2 |
|  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 |
|  | Explicit VR Big Endian | 1.2.840.10008.1.2.2 |
|  | JPEG Baseline (Process 1) | 1.2.840.10008.1.2.4.50 |
|  | JPEG Lossless (Process 14) Selection Value 1 | 1.2.840.10008.1.2.4.70 |
|  | RLE Lossless | 1.2.840.10008.1.2.5 |
|  | Other (please specify): |  |

General Comments on the Proposed Transfer Syntax Table:

Comments

### Other Association Negotiation Items

| Yes | Application Context |  |
| --- | --- | --- |
|  | DICOM Standard Application Context | 1.2.840.10008.3.1.1.1 |

If the application supports an asynchronous operation, please specify the Maximum-number-operations-invoked and Maximum-number-operations-performed values sent during association negotiation.

Asynchronous Operations Window

If the application supports a SCU/SCP role selection, please specify which roles are proposed for which SOP class.

SCU/SCP Role Selection

If the application supports extended negotiation for any of the proposed SOP classes, please specify.

Extended Negotiation

General Comments on the Association Negotiation Items:

Comments

## Data Structure transmitted with the Storage Service SOP Classes

In this section you have to specify the attributes which may be contained in the Image IOD sent with a STORE message.  All attributes which are relevant for the image display are listened.

### Modality LUT Transformation

Please specify whether the Image IOD sent with a STORE message contains a Modality LUT transformation.  All possible combinations of the three LUT descriptor values should be specified.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Modality LUT Sequence | (0028,3000) |  |
|  | >LUT Descriptor | (0028,3002) |  |
|  | >LUT Explanation | (0028,3003) |  |
|  | >Modality LUT Type | (0028,3004) |  |
|  | >LUT Data | (0028,3006) |  |
|  | Rescale Intercept | (0028,1052) |  |
|  | Rescale Slope | (0028,1053) |  |
|  | Rescale Type | (0028,1054) |  |

### Mask Transformation

Please specify whether the Image IOD sent with a STORE message contains a Mask transformation.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Mask Subtraction Sequence | (0028,6010) |  |
|  | >Mask Operation | (0028,6101) |  |
|  | >Applicable Frame Range | (0028,6102) |  |
|  | >Mask Frame Numbers | (0028,6110) |  |
|  | >Contrast Frame Averaging | (0028,6112) |  |
|  | >Mask Sub-pixel Shift | (0028,6114) |  |
|  | >TID Offset | (0028,6120) |  |
|  | >Mask Operations Explanations | (0028,6190) |  |
|  | Recommended Viewing Mode | (0028,1090) |  |

### VOI LUT Transformation

Please specify whether the Image IOD sended with a STORE message contains a VOI LUT transformation.  All possible combinations of the three LUT descriptor values should be specified.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | VOI LUT Sequence | (0028,3010) |  |
|  | >LUT Descriptor | (0028,3002) |  |
|  | >LUT Explanation | (0028,3003) |  |
|  | >LUT Data | (0028,3006) |  |
|  | Window Center | (0028,1050) |  |
|  | Rescale Width | (0028,1051) |  |
|  | Window Center & Width Explanation | (0028,1055) |  |

### Presentation LUT Transformation

Please specify whether the Image IOD sent with a STORE message contains a Presentation LUT transformation.  All possible combinations of the three LUT descriptor values should be specified.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Presentation LUT Shape | (2050,0020) |  |

### Overlay

Please specify whether the Image IOD sent with a STORE message contains an overlay.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Overlay Rows | (60xx,0010) |  |
|  | Overlay Columns | (60xx,0011) |  |
|  | Overlay Type | (60xx,0040) |  |
|  | Overlay Origin | (60xx,0050) |  |
|  | Overlay Bits Allocated | (60xx,0100) |  |
|  | Overlay Bit Position | (60xx,0102) |  |
|  | Overlay Data | (60xx,3000) |  |
|  | Overlay Description | (60xx,0022) |  |
|  | Overlay Subtype | (60xx,0045) |  |
|  | Overlay Label | (60xx,1500) |  |
|  | ROI Area | (60xx,1301) |  |
|  | ROI Mean | (60xx,1302) |  |
|  | ROI Standard Deviation | (60xx,1303) |  |
|  | Attributes not included in the Overlay Plane Module |  |  |
|  | Number of Frames in Overlay | (60xx,0015) |  |
|  | Image Frame Origin | (60xx,0051) |  |
|  | Overlay Planes | (60xx,0012) |  |
|  | Overlay Planes Origin | (60xx,0052) |  |

General Comments

Comments

### Curve

Please specify whether the Image IOD sent with a STORE message contains a curve.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Curve Dimensions | (50xx,0005) |  |
|  | Number of Points | (50xx,0010) |  |
|  | Type of Data | (50xx,0020) |  |
|  | Data Value Representation | (50xx,0103) |  |
|  | Curve Data | (50xx,3000) |  |
|  | Curve Description | (50xx,0022) |  |
|  | Axis Units | (50xx,0030) |  |
|  | Axis Labels | (50xx,0040) |  |
|  | Minimum Coordinate Value | (50xx,0104) |  |
|  | Maximum Coordinate Value | (50xx,0105) |  |
|  | Curve Range | (50xx,0106) |  |
|  | Curve Data Descriptor | (50xx,0110) |  |
|  | Coordinate Start Value | (50xx,0112) |  |
|  | Coordinate Step Value | (50xx,0114) |  |
|  | Curve Label | (50xx,2500) |  |
|  | Referenced Overlay Sequence | (50xx,2600) |  |
|  | >Referenced SOP Class UID | (0008,1150) |  |
|  | >Referenced SOP Instance UID | (0008,1155) |  |
|  | >Referenced Overlay Group | (50xx,2610) |  |

### Image Pixel related Attributes

Please specify which attributes and attribute values may be sent as part of an image IOD with a STORE message.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Samples Per Pixel | (0028,0002) |  |
|  | Photometric Interpretation | (0028,0004) |  |
|  | Rows | (0028,0010) |  |
|  | Columns | (0028,0011) |  |
|  | Bits Allocated | (0028,0100) |  |
|  | Bits Stored | (0028,0101) |  |
|  | High Bit | (0028,0102) |  |
|  | Pixel Representation | (0028,0103) |  |
|  | Pixel Data | (7FE0,0010) |  |
|  | Planar Configuration | (0028,0006) |  |
|  | Pixel Aspect Ratio | (0028,0034) |  |
|  | Smallest Image Pixel Value | (0028,0106) |  |
|  | Largest Image Pixel Value | (0028,0107) |  |
|  | Red Palette Color Lookup Table Descriptor | (0028,1101) |  |
|  | Green Palette Color Lookup Table Descriptor | (0028,1102) |  |
|  | Blue Palette Color Lookup Table Descriptor | (0028,1103) |  |
|  | Red Palette Color Lookup Table Data | (0028,1201) |  |
|  | Green Palette Color Lookup Table Data | (0028,1202) |  |
|  | Blue Palette Color Lookup Table Data | (0028,1203) |  |
|  | Attributes not included in the Image Pixel Module: |  |  |
|  | Shutter Shape | (0018,1600) |  |
|  | Pixel Spacing | (0028,0030) |  |
|  | Number of Frames | (0028,0008) |  |
|  | Frame Increment Pointer | (0028,0009) |  |
|  | Planes | (0028,0012) |  |
|  | Other (please specify): |  |  |

General Comments

Comments

## Data Structure transmitted with the Grayscale Softcopy Presentation State Storage SOP Class

In this section you have to specify the attributes which may be contained in the Presentation State IOD sent with a STORE message.  All attributes which are relevant for the image display and are listed.

### Modality LUT Transformation

Please specify whether the Presentation State IOD sent with a STORE message contains a modality transformation.  All possible combinations of the three LUT descriptor values should be specified.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Modality LUT Sequence | (0028,3000) |  |
|  | >LUT Descriptor | (0028,3002) |  |
|  | >LUT Explanation | (0028,3003) |  |
|  | >Modality LUT Type | (0028,3004) |  |
|  | >LUT Data | (0028,3006) |  |
|  | Rescale Intercept | (0028,1052) |  |
|  | Rescale Slope | (0028,1053) |  |
|  | Rescale Type | (0028,1054) |  |

### Mask Transformation

Please specify whether the Presentation State IOD sent with a STORE message contains a mask transformation

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Mask Subtraction Sequence | (0028,6010) |  |
|  | >Mask Operation | (0028,6101) |  |
|  | >Applicable Frame Range | (0028,6102) |  |
|  | >Mask Frame Numbers | (0028,6110) |  |
|  | >Contrast Frame Averaging | (0028,6112) |  |
|  | >Mask Sub-pixel Shift | (0028,6114) |  |
|  | >TID Offset | (0028,6120) |  |
|  | >Mask Operations Explanations | (0028,6190) |  |
|  | Recommended Viewing Mode | (0028,1090) |  |

### VOI LUT Transformation

Please specify whether the Presentation State IOD sent with a STORE message contains a VOI transformation.  All possible combinations of the three LUT descriptor values should be specified.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Softcopy VOI LUT Sequence | (0028,3010) |  |
|  | >Referenced Image Sequence | (0008,1140) |  |
|  | >>Referenced SOP Class UID | (0008,1150) |  |
|  | >>Referenced SOP Instance UID | (0008,1155) |  |
|  | >>Referenced Frame Number | (0008,1160) |  |
|  | >VOI LUT Sequence | (0028,3110) |  |
|  | >>LUT Descriptor | (0028,3002) |  |
|  | >>LUT Explanation | (0028,3003) |  |
|  | >>LUT Data | (0028,3006) |  |
|  | >Window Center | (0028,1050) |  |
|  | >Rescale Width | (0028,1051) |  |
|  | >Window Center & Width Explanation | (0028,1055) |  |

### Presentation LUT Transformation

Please specify whether the Presentation State IOD sent with a STORE message contains a presentation LUT transformation.  All possible combinations of the three LUT descriptor values should be specified.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Presentation LUT Sequence | (2050,0010) |  |
|  | >LUT Descriptor | (0028,3002) |  |
|  | >LUT Explanation | (0028,3003) |  |
|  | >LUT Data | (0028,3006) |  |
|  | Presentation LUT Shape | (2050,0020) |  |

### Shutter Transformation

Please specify whether the Presentation State IOD sent with a STORE message contains shutters.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Shutter Shape | (0018,1600) |  |
|  | Shutter Presentation Value | (0018,1622) |  |
|  | Shutter Overlay Group | (0018,1623) |  |
|  | Shutter Overlay Group | (0018,1623) |  |

### Image Annotation

Please specify attributes related to the image annotations.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Graphic Annotation Sequence | (0070,0001) |  |
|  | >Referenced Image Sequence | (0008,1140) |  |
|  | >>Referenced SOP Class UID | (0008,1150) |  |
|  | >>Referenced SOP Instance UID | (0008,1155) |  |
|  | >>Referenced Frame Number | (0008,1160) |  |
|  | >Graphic Layer | (0070,0002) |  |
|  | >Text Object Sequence | (0070,0008) |  |
|  | >>Bounding Box Annotation Units | (0070,0003) |  |
|  | >>Anchor Point Annotation Units | (0070,0004) |  |
|  | >>Unformatted Text Value | (0070,0006) |  |
|  | >>Bounding Box Top Left Hand Corner | (0070,0010) |  |
|  | >>Bounding Box Bottom Right Hand Corner | (0070,0011) |  |
|  | >>Bounding Box Text Horizontal Justification | (0070,0012) |  |
|  | >>Anchor Point | (0070,0014) |  |
|  | >>Anchor Point Visibility | (0070,0015) |  |
|  | >Graphic Object Sequence | (0070,0009) |  |
|  | >>Graphic Annotation Units | (0070,0005) |  |
|  | >>Graphic Dimensions | (0070,0020) |  |
|  | >>Number of Graphic Points | (0070,0021) |  |
|  | >> Graphic Data | (0070,0022) |  |
|  | >>Graphic Type | (0070,0023) |  |
|  | >>Graphic Filled | (0070,0024) |  |

### Spatial Transformation

Please specify whether the Presentation State IOD sent with a STORE message contains spatial transformations.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Image Rotation | (0070,0042) |  |
|  | Image Horizontal Flip | (0070,0041) |  |

### Overlay

Please specify whether the Presentation State IOD sent with a STORE message contains an overlay.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Overlay Rows | (60xx,0010) |  |
|  | Overlay Columns | (60xx,0011) |  |
|  | Overlay Type | (60xx,0040) |  |
|  | Overlay Origin | (60xx,0050) |  |
|  | Overlay Bits Allocated | (60xx,0100) |  |
|  | Overlay Bit Position | (60xx,0102) |  |
|  | Overlay Data | (60xx,3000) |  |
|  | Overlay Description | (60xx,0022) |  |
|  | Overlay Subtype | (60xx,0045) |  |
|  | Overlay Label | (60xx,1500) |  |
|  | ROI Area | (60xx,1301) |  |
|  | ROI Mean | (60xx,1302) |  |
|  | ROI Standard Deviation | (60xx,1303) |  |

General Comments

Comments

### Curve

Please specify whether the Presentation State IOD sent with a STORE message contains a curve.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Curve Dimensions | (50xx,0005) |  |
|  | Number of Points | (50xx,0010) |  |
|  | Type of Data | (50xx,0020) |  |
|  | Data Value Representation | (50xx,0103) |  |
|  | Curve Data | (50xx,3000) |  |
|  | Curve Description | (50xx,0022) |  |
|  | Axis Units | (50xx,0030) |  |
|  | Axis Labels | (50xx,0040) |  |
|  | Minimum Coordinate Value | (50xx,0104) |  |
|  | Maximum Coordinate Value | (50xx,0105) |  |
|  | Curve Range | (50xx,0106) |  |
|  | Curve Data Descriptor | (50xx,0110) |  |
|  | Coordinate Start Value | (50xx,0112) |  |
|  | Coordinate Step Value | (50xx,0114) |  |
|  | Curve Label | (50xx,2500) |  |
|  | Referenced Overlay Sequence | (50xx,2600) |  |
|  | >Referenced SOP Class UID | (0008,1150) |  |
|  | >Referenced SOP Instance UID | (0008,1155) |  |
|  | >Referenced Overlay Group | (50xx,2610) |  |

### Displayed Area

Please specify the attributes related to the displayed area.

| Yes | Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- | --- |
|  | Displayed Area Selection Sequence | (0070,005A) |  |
|  | >Referenced Image Sequence | (0008,1140) |  |
|  | >>Referenced SOP Class UID | (0008,1150) |  |
|  | >>Referenced SOP Instance UID | (0008,1155) |  |
|  | >>Referenced Frame Number | (0008,1160) |  |
|  | >Displayed Area Top Left Hand Corner | (0070,0052) |  |
|  | >Displayed Area Bottom Right Hand Corner | (0070,0053) |  |
|  | >Presentation Size Mode | (0070,0100) |  |
|  | >Presentation Pixel Spacing | (0070,0101) |  |
|  | >Presentation Pixel Aspect Ratio | (0070,0102) |  |
|  | >Presentation Pixel Magnification Ratio | (0070,0103) |  |

### Other Attributes

| Attribute Name | Tag | Attribute Value(s) |
| --- | --- | --- |

## Behavior of other SOP Classes

Please describe the behavior of other proposed SOP classes.

Behavior of other Proposed SOP Classes

# Test Procedure for Image Displays

The IHE Year 2 MESA Tools contain, among other components, a Java-based GUI application called “DICOMscope”.  This application is used in the test procedure for Image Creators.  Installation instructions and a simple user manual for the DICOMscope application are provided in different documents with the IHE Year 2 MESA distribution.  The following discussion assumes that the application has been successfully installed and that the user has a basic understanding of the application’s concepts as laid out in the user manual.  However, not all functions offered by the DICOMscope application are required or used in this test procedure.

DICOMscope contains a Query/Retrieve Server that implements the requirements of the IHE Year 2 “Query Image”, “Retrieve Image”, “Query Presentation State”, “Retrieve Presentation State” transactions.  It should be noted that interoperability between an Image Composer and the DICOMscope testing Query/Retrieve Server does improve the probability for interoperability between the SCU and SCP of the specified provided by a different vendor, but is in no way a guarantee for interoperability.

Images and Presentation States created with any Image Creator should be validated with DICOMscope.  DICOMscope is a Storage Service Class Provider.  Every Image Creator can send images and Presentation States to DICOMscope where they can be visualized.  It should be noted that DICOMscope is not able to display color images and that Presentation States are only defined for grayscale images.

The Image Creator should be grouped with an Image Creator.  Therefore, the Test Plan for the Image Display should be checked.