HIMMS and RSNA

IHE

Integrating the Healthcare Enterprise

Test Plan for Image Displays

2003-12-19

Marco Eichelberg1,  Klaus Kleber2, Jörg Riesmeier1

1	Kuratorium OFFIS e.V., Escherweg 2, 26121 Oldenburg, Germany

2	VISUS Technology Transfer GmbH, Philippstraße 5, 44803 Bochum, Germany

Contents

1	Introduction	4

1.1	Acknowledgements	4

2	Test Procedure for Image Displays	5

2.1	Prerequisites	5

2.2	Testing the GSDF Calibration of an Image Display Actor	5

2.2.1	Open Excel File and Fill Administrative Data	6

2.2.2	Load Images, Measure Boxes and Enter Results	6

2.2.3	Print “Output” Page and Assess Deviation from GSDF	7

2.3	Testing Image Display without Presentation States	8

2.3.1	Modality LUT	8

2.3.2	VOI LUT	11

2.3.3	Modality / VOI LUTs	13

2.4	Testing Image Display with Presentation States	14

2.4.1	Modality LUT	14

2.4.2	VOI LUT	17

2.4.3	Presentation LUT	19

2.4.4	Modality / VOI / Presentation LUTs	22

2.4.5	Spatial Transformation	22

2.4.6	Displayed Area	25

2.4.7	Graphic Annotation	26

2.4.8	Text Annotation	31

2.4.9	Display Shutter	35

2.4.10	Overlays	37

2.4.11	Complex Combination	38

3	Image Display Test Results	41

3.1	Image Display without Presentation States	41

3.1.1	Modality LUT	41

3.1.2	VOI LUT	42

3.1.3	Modality / VOI LUTs	42

3.2	Image Display with Presentation States	42

3.2.1	Modality LUT	43

3.2.2	VOI LUT	43

3.2.3	Presentation LUT	44

3.2.4	Modality / VOI / Presentation LUTs	44

3.2.5	Spatial Transformations	44

3.2.6	Displayed Area	45

3.2.7	Graphic Annotation	45

3.2.8	Text Annotation	46

3.2.9	Display Shutter	46

3.2.10	Overlay	47

3.2.11	Complex Combination	47

3.3	General Comments	47

4	References	48

# Introduction

This document is intended for those vendors participating in the IHE demonstration which will supply one or more Image Display Actors as defined in the IHE Radiology Technical Framework. This document specifies a number of test cases which allow to assess:

- whether the Image Display is correctly calibrated to the DICOM Grayscale Standard Display Function (GSDF) [1] as required by the Technical Framework,

- whether the display of a set of test images (without presentation states) leads to the expected results, and

- whether the display of a set of test images with presentation states leads to the expected results.

Section 2 describes prerequisites of this test plan, how to set up and perform the test cases, and how to use the Excel spreadsheet that is used to assess the measured results from many of the test cases.

Section 3 is a form to be filled in with the Image Display’s test results.  If an Image Display application can be operated in different modes which exhibit different behavior, the tests should be run for each mode of operation, and multiple copies of the form should be filled in.

## Acknowledgements

We would like to thank David Clunie (dclunie@dclunie.com) for preparing and providing a set of test images and presentation states.   These DICOM images and presentation states were used to create most test cases in section 3.4.  The original package can be downloaded from ftp://medical.nema.org/medical/dicom/rsnademos/display/.

Since its initial release in 2000, many people have contributed to this test suite by reporting bugs or even providing bug fixes and new test cases (in alphabetical order):

- Jan-Pieter Diender, GEMS IT Applicare (reported bug in “spat” cases)

- Dave Harvey, Medical Connections (reported bug in “disa” cases and small inconsistencies in various test cases)

- Sebastian Hauer, TIANI MedGraph (provided new “spat” cases)

- Gunter Zeilinger, TIANI MedGraph (reported bug in “mlut” cases)

- Dave Maffit, Steve Moore and Eric Poiseau from IHE

... and some anonymous IHE participants. Thank you.

# Test Procedure for Image Displays

## Prerequisites

The following items are required for performance of the Test Plan:

- Photometer for measuring the luminance of a softcopy display device and the luminance contribution from reflected ambient light.  Suitable devices for this purpose are identified in a separate document entitled “Calibration Tools”.  It should be noted that this Test Plan is based on the international system of units (SI).  Therefore, luminance measurements in candela per square meter (cd/m²) are required.

- Image Display Test Cases consisting of a number of images and presentation states that will be displayed during the execution of the test plan.  These are provided as part of the IHE MESA tools distribution.  Note that a small gray ramp has been added to the lower border of most images. This prevents applications from masquerading the test results by computing an automatic VOI window.

- Microsoft Excel 97 Spreadsheet “gsdf_lum.xls”. This spreadsheet is explained in section 3.4.  Also provided as part of the IHE MESA tools distribution.

## Testing the GSDF Calibration of an Image Display Actor

The Image Display Test Cases contain a set of 17 DICOM secondary capture images (512*512 pixel resolution, 8 bits/pixel depth) intended for an assessment of the correct GSDF calibration of a display system.  These images are part of the Image Display Test Cases and can be identified by the filenames “ddl_01.dcm” to “ddl_17.dcm”.  Each image shows a single box of gray as described in DICOM Part 14 [1]: the measurement field is centered and has a size of 10% of the total number of image pixels (162*162), the background is uniform with a luminance of 20% of the maximum luminance (corresponds to 50% gray).  Figure 1 shows one of the test images – the image number referenced by the Excel form (“measured DDL #”) is also displayed in the top left-hand corner.

Figure 1: Test image with measurement field

The Excel file “gsdf_lum.xls” contains two tables: the first table is used to enter the measured luminance values and some administrative data, the second one is automatically filled and used for printing the results.  Figure 2 shows the “Input” table – the sheet is protected and only the blue fields can be modified.

Figure 2: “Input” table of the Excel file for softcopy calibration

Note: The Excel file contains macros which are required for perfect operation, i. e. they should be activated if the user is asked by the application.

### Open Excel File and Fill Administrative Data

First you should open the Excel file “gsdf_lum.xls” (activate macros if a requesting dialog appears) and enter the administrative data: Vendor / Device, Date / Time, Performed by.

### Load Images, Measure Boxes and Enter Results

Start the DICOM viewer application you want to check for correct GSDF implementation and load the first test image “ddl_01.dcm”.  It is very important that the full range of 256 shades of gray is used to display the image.  Therefore, the VOI window 128/256 (center/width) is stored in the file (assuming that the VOI LUT transformation has been implemented according to the DICOM standard, PS3.3 [2] section C.11).  Additionally, all 256 shades of gray are used in the gray ramp at the lower end of the image to support viewers which are ignoring the stored settings and initially using an automatically computed window.  However, if the viewer does not use the full range of grayscale values, please, make sure that you change the VOI window manually before measuring.

In order to correspond with the instructions contained in DICOM Part 14 the image should be scaled to fit the full screen and the measurement field (gray box) should be on the center of the screen.  Then the luminance meter should be used according to the description shipped with the device.  The measured value (in cd/m²) should be entered into the corresponding field in the Excel “Input” table (e. g. “measured DDL #1” for the first image “ddl_001.dcm”).  Please also compare the number displayed in the top left-hand corner to avoid mistakes.

The same procedure has to be performed for the other 16 test images.  After all 17 measured values have been entered the “overall Status” should be “OK”.  Otherwise you have most likely made a mistake.

Depending on the calibration tool the ambient light is measured separately.  In this case you need to enter the measured ambient light in the field below the administrative data.  Please note that the measurement unit is also candela per square meter.

### Print “Output” Page and Assess Deviation from GSDF

After successfully measuring and entering the data the „Output“ page shows the result (see Figure 3).

This page can be printed on standard paper.  Since all necessary information is contained on this page there is no need to print the “Input” table.

The table below the administrative data shows the measured luminance values versus the GSDF and the deviation from the standard curve (absolute and percentage values as well as the mean difference and the standard deviation).  Unfortunately, no precise formula for an assessment of the deviation can be given.  Obviously, deviation from the GSDF should be as small as possible and in any case below 20%.  In addition, the curve should be “smooth” and should not show any steps.  This is the reason why both curves are shown in a diagram (where the luminance is logarithmically scaled).  The 17 measured values are interpolated by a built-in algorithm – though, this has no influence on the original data presented in the table.

Figure 3: Printed “Output” table of the Excel file for softcopy calibration

## Testing Image Display without Presentation States

This section describes a set of tests for Image Display without Presentation States.  Each test consists of an image (“*.dcm”) which should be displayed by the Image Display Actor, a description of the relevant attributes in the Image IOD and a description of the correct resulting display for evaluation purposes.

### Modality LUT

This subsection contains tests for the Modality LUT transformation.

| Test MLUT_01 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation MONOCHROME2, Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_02 has been removed.

| Test MLUT_03 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range 0 to +4095, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept -128 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_04 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range 0 to +4095, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept -1024 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_05 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, high bits 1, Pixel Range 0 to +4095, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_06 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, high bits 1, Pixel Range 0 to +4095, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept -1024 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_07 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 15, Pixel Range 0 to +32767, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_08 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 15, Pixel Range 0 to +32767, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept -1024 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_09 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range 0 to +65535, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_10 has been removed.

| Test MLUT_11 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_12 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept -1024 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_13 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range -32768 to +32767, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTS |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_14 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range -32768 to +32767, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_15 has been removed.

| Test MLUT_16 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range -32768 to +32767, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, Rescale Slope 0.5, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_17 has been removed.

| Test MLUT_18 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, Modality LUT Sequence with LUT Descriptor 4096\63488\16, linear LUT value o up to 65535 with increment +16 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_19 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME1 (inverted), Modality LUT Sequence with LUT Descriptor 4096\63488\16, linear LUT value 65535 down to 0 with increment -16 |
| Result | Display should be a normal SMPTE pattern |

### VOI LUT

This subsection contains tests for the VOI LUT transformation.  Though presentation states are not part of the tests in this section it is expected that the VOI LUT transformation is implemented according to the DICOM standard, PS3.3 [2] section C.11.

| Test VLUT_01 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2 , no LUTs |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_02 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Window Center = 2^(Bits Stored-1), Window Width = 2^Bits Stored, i. e. 128/256, i. e. identity |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_03 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range +25 to +75, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Window Center = 50.5, Window Width = 51 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_04 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, VOI LUT Sequence with LUT Descriptor 256\0\16,linear LUT value 0 up to 65535 with increment +257 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_05 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME1 (inverted), VOI LUT Sequence with LUT Descriptor 256\0\16, linear LUT value 65535 down to 0 with increment -257 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_06 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_07 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs; Window Center = 0, Window Width = 4096, i. e. identity |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_08 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range 25 to +74, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, Window Center = 50, Window Width = 50 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_09 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, VOI LUT Sequence with LUT Descriptor 4096\63488\16, linear LUT value 0 up to 65535 with increment +16 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_10 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, VOI LUT Sequence with LUT Descriptor 4096\06344\16, linear LUT value 65535 down to 0 with increment -16 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_11 |  |
| --- | --- |
| Image | SC image with two gray ramps (with 256/16 shades of gray), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +256, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, 9 different Window Center and Width settings (128/256, 128/1, 128/16, 64/16, 128/32, 32/64, 218/96, 48/96, 128/128) |
| Result | Lower part of the image should display as many different gray boxes as indicated in the corresponding Window Center and Width Explanation (full range, black and white, 2 gray boxes, 3 gray boxes, 4 gray boxes, 5 gray boxes, 6 gray boxes, 7 gray boxes, 10 gray boxes). NB: If the grayscale correction according to the GSDF is done “in software” it might happen that the number of visible gray boxes differs from that specified above. When using DICOMscope the corection should be switched off in this case. |

### Modality / VOI LUTs

This subsection contains combined tests for Modality and VOI LUT.

| Test XLUT_01 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 128 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 2, Rescale Intercept -128, Window Center = 255.5, Window Width = 255 |
| Result | Display should be a normal SMPTE pattern |

| Test XLUT_02 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range 1024 to +65535, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, Rescale Slope 0.125, Rescale Intercept -1024, Window Center = 3136, Window Width = 8064 |
| Result | Display should be a normal SMPTE pattern |

## Testing Image Display with Presentation States

This section describes a set of tests for Image Display with Presentation States.  Each test consists of an Image (“*.dcm”) and a Presentation State (“*.pre”) which should be displayed by the Image Display Actor (i. e. the presentation state should be applied to the image).  For both a description of the relevant attributes is listed as well as a description of the correct resulting display for evaluation purposes.

For some of the test cases a result image (“*.res”) is provided in order to facilitate the decision whether an implementation is correct or not.  However, this does not imply that the rendered output has to match exactly to be standard compliant (deviations might result from different fonts, divergent visualization of graphic annotations, etc.)

### Modality LUT

This subsection contains tests for the Modality LUT transformation.

| Test MLUT_P01 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_P02 has been removed.

| Test MLUT_P03 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range 0 to +4095, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P04 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range 0 to +4095, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept –1024 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P05 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, high bits 1, Pixel Range 0 to +4095, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P06 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, high bits 1, Pixel Range 0 to +4095, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept -1024 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P07 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 15, Pixel Range 0 to +32767, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P08 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 15, Pixel Range 0 to +32767, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept -1024 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P09 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range 0 to +65535, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_P10 has been removed.

| Test MLUT_P11 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P12 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept -1024 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P13 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range -32768 to +32767, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | No LUTS |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P14 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range -32768 to +32767, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 1, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_P15 has been removed.

| Test MLUT_P16 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 16, Pixel Range -32768 to +32767, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Rescale Slope 0.5, Rescale Intercept 0 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_P17 has been removed.

| Test MLUT_P18 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Modality LUT Sequence with LUT Descriptor 4096\63488\16, linear LUT value o up to 65535 with increment +16 |
| Result | Display should be a normal SMPTE pattern |

| Test MLUT_P19 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME1 (inverted), no LUTs |
| PState | Modality LUT Sequence with LUT Descriptor 4096\63488\16, linear LUT value 65535 down to 0 with increment -16 |
| Result | Display should be a normal SMPTE pattern |

Test MLUT_P20 has been removed.

Test MLUT_P21 has been removed.

### VOI LUT

This subsection contains tests for the VOI LUT transformation.

| Test VLUT_P01 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | No LUT |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P02 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Window Center = 2^(Bits Stored-1), Window Width = 2^Bits Stored, i. e. 128/256, i. e. identity |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P03 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range +25 to +75, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Window Center = 50.5, Window Width = 51 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P04 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | VOI LUT Sequence with LUT Descriptor 256\0\16,linear LUT value 0 up to 65535 with increment +257 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P05 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME1 (inverted), no LUTs |
| PState | VOI LUT Sequence with LUT Descriptor 256\0\16, linear LUT value 65535 down to 0 with increment -257 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P06 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | absent ... default VOI LUT to identity |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P07 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Window Center = 0, Window Width = 4096, i. e. identity |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P08 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range 25 to +74, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Window Center = 50, Window Width = 50 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P09 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | VOI LUT Sequence with LUT Descriptor 4096\63488\16, linear LUT value 0 up to 65535 with increment +16 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P10 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | VOI LUT Sequence with LUT Descriptor 4096\63488\16, linear LUT value 65535 down to 0 with increment -16 |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P11 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range +25 to +75, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, inappropriate VOI LUT transform (Window Center 128, Window Width 256) |
| PState | Appropriate VOI LUT transform (Window Center = 50.5, Window Width = 51) |
| Result | Display should be a normal SMPTE pattern |

| Test VLUT_P12 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to 255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, inappropriate VOI LUT transform (Window Center 0, Window Width 128), requiring identity VOI LUT transform |
| PState | VOI LUT transform not specified so assumed identity |
| Result | Display should be a normal SMPTE pattern |

### Presentation LUT

This subsection contains tests for the Presentation LUT transformation.

| Test PLUT_P01 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation LUT Shape = IDENTITY |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P02 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation LUT Shape = INVERSE |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P03 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation LUT Shape = IDENTITY |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P04 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME1 (inverted), no LUTs |
| PState | Presentation LUT Shape = INVERSE |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P05 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation LUT Sequence with LUT Descriptor 256\0\10, linear LUT value 0 up to 1023 with increment +4 |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P06 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME1 (inverted), no LUTs |
| PState | Presentation LUT Sequence with LUT Descriptor 256\0\10, linear LUT value 1023 down to 0 with increment -4 |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P07 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation LUT Sequence with LUT Descriptor 4096\63488\16, linear LUT value 0 up to 65535 with increment +16 |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P08 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation LUT Sequence with LUT Descriptor 4096\63488\8, linear LUT value 0 up to 255 |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P09 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, inappropriate Presentation LUT transform (Presentation LUT Shape = INVERSE) |
| PState | Appropriate Presentation LUT transform (Presentation LUT Shape = IDENTITY) |
| Result | Display should be a normal SMPTE pattern |

| Test PLUT_P10 |  |
| --- | --- |
| Image | SC image with inverse SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation LUT Sequence with LUT Descriptor 256\0\8 using 16 bits per LUT entry (512 byte LUT data), linear LUT value 255 down to 0 |
| Result | Display should be a normal SMPTE pattern |

### Modality / VOI / Presentation LUTs

This subsection contains combined tests for the different LUTs (Modality LUT, VOI LUT, Presentation LUT) in the DICOM transformation pipeline.

Test XLUT_P01 has been removed.

| Test XLUT_P02 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 16, Bits Stored 12, Pixel Range -2048 to +2047, Pixel Representation is signed, Photometric Interpretation is monochrom2 , no LUTs |
| PState | Rescale Slope -2, Rescale Intercept -1; Window Center 0, Window Width 8192; Presentation LUT Shape INVERSE |
| Result | Display should be a normal SMPTE pattern |

| Test XLUT_P03 |  |
| --- | --- |
| Image | SC image with two gray ramps (with 256/16 shades of gray), pixel data scrambled by the inverse of the three LUTs stored in the presentation state |
| PState | Modality LUT Sequence with LUT Descriptor 256\0\8, scrambled LUT data; VOI LUT Sequence with LUT Descriptor 256\0\8, scrambled LUT data; Presentation LUT Sequence with LUT Descriptor 256\0\8, scrambled LUT data |
| Result | Result on display should look like “xlut_p03.res” |

### Spatial Transformation

This subsection contains tests for the following spatial transformations: rotation and horizontal flip.

| Test SPAT_P01 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Image Rotation 0, Image Horizontal Flip N |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P02 |  |
| --- | --- |
| Image | SC image with SMPTE pattern (rotation 270 degree CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Image Rotation 90, Image Horizontal Flip N |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P03 |  |
| --- | --- |
| Image | SC image with SMPTE pattern (rotation 180 degree CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Image Rotation 180, Image Horizontal Flip N |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P04 |  |
| --- | --- |
| Image | SC image with SMPTE pattern (rotation 90 degree CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Image Rotation 270, Image Horizontal Flip N |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P05 |  |
| --- | --- |
| Image | SC image with SMPTE pattern (horizontal fliped), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Image Rotation 0, Image Horizontal Flip Y |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P06 |  |
| --- | --- |
| Image | SC image with SMPTE pattern (horizontal fliped and rotation 270 degrees CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Image Rotation 90, Image Horizontal Flip Y |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P07 |  |
| --- | --- |
| Image | SC image with SMPTE pattern (horizontal fliped and rotation 180 degrees CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Image Rotation 180, Image Horizontal Flip Y |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P08 |  |
| --- | --- |
| Image | SC image with SMPTE pattern (horizontal fliped and rotation 90 degrees CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Image Rotation 270, Image Horizontal Flip Y |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P09 |  |
| --- | --- |
| Image | SC image (Columns 1024, Rows 1024) with 512*512 SMPTE pattern (rotation 270 degree CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1024, Displayed Area Bottom Right Hand Corner 512\513, Presentation Pixel Aspect Ratio 1\1, Image Rotation 90, Image Horizontal Flip N |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P10 |  |
| --- | --- |
| Image | SC image (Columns 1024, Rows 1024) with 512*512 SMPTE pattern (rotation 180 degree CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 512\1024, Displayed Area Bottom Right Hand Corner 1\513, Presentation Pixel Aspect Ratio 1\1, Image Rotation 180, Image Horizontal Flip N |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P11 |  |
| --- | --- |
| Image | SC image (Columns 1024, Rows 1024) with 512*512 SMPTE pattern (rotation 90 degree CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 512\513, Displayed Area Bottom Right Hand Corner 1\1024, Presentation Pixel Aspect Ratio 1\1, Image Rotation 270, Image Horizontal Flip N |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P12 |  |
| --- | --- |
| Image | SC image (Columns 1024, Rows 1024) with 512*512 SMPTE pattern (horizontal fliped), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 512\513, Displayed Area Bottom Right Hand Corner 1\1024, Presentation Pixel Aspect Ratio 1\1, Image Rotation 0, Image Horizontal Flip Y |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P13 |  |
| --- | --- |
| Image | SC image (Columns 1024, Rows 1024) with 512*512 SMPTE pattern (horizontal fliped and rotation 270 degrees CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\513, Displayed Area Bottom Right Hand Corner 512\1024, Presentation Pixel Aspect Ratio 1\1, Image Rotation 90, Image Horizontal Flip Y |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P14 |  |
| --- | --- |
| Image | SC image (Columns 1024, Rows 1024) with 512*512 SMPTE pattern (horizontal fliped and rotation 180 degrees CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1024, Displayed Area Bottom Right Hand Corner 512\513, Presentation Pixel Aspect Ratio 1\1, Image Rotation 180, Image Horizontal Flip Y |
| Result | Display should be a normally oriented SMPTE pattern |

| Test SPAT_P15 |  |
| --- | --- |
| Image | SC image (Columns 1024, Rows 1024) with 512*512 SMPTE pattern (horizontal fliped and rotation 90 degrees CW), Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 512\1024, Displayed Area Bottom Right Hand Corner 1\513, Presentation Pixel Aspect Ratio 1\1, Image Rotation 270, Image Horizontal Flip Y |
| Result | Display should be a normally oriented SMPTE pattern |

### Displayed Area

This subsection contains tests for the displayed area.

| Test DISA_P01 |  |
| --- | --- |
| Image | SC image containing SMTPE pattern (centered), Columns 1280, Rows 900, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 385\195, Displayed Area Bottom Right Hand Corner 896\706, Presentation Pixel Aspect Ratio 1\1 |
| Result | Display should be a normal SMPTE pattern (full screen) |

| Test DISA_P02 |  |
| --- | --- |
| Image | SC image containing SMTPE pattern (placed in bottom right-hand corner), Columns 1280, Rows 900, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 769\389, Displayed Area Bottom Right Hand Corner 1280\900, Presentation Pixel Aspect Ratio 1\1 |
| Result | Display should be a normal SMPTE pattern (full screen) |

| Test DISA_P03 |  |
| --- | --- |
| Image | SC image containing SMTPE pattern (magnified by factor 2.0), Columns 1024, Rows 1024, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: MAGNIFY, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 1024\1024, Presentation Pixel Aspect Ratio 1\1, Presentation Pixel Magnification Ratio 0.5 |
| Result | Display should be a normal SMPTE pattern (fixed size: 512*512 pixel) |

| Test DISA_P04 |  |
| --- | --- |
| Image | SC image containing SMTPE pattern (magnified by factor 2.0 and 3.0 respectively), Columns 1024, Rows 1536, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: TRUE SIZE, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 1024\1536, Presentation Pixel Spacing 0.1302083\0.1953125 |
| Result | Display should be a normal SMPTE pattern (size: 200*200 mm) |

| Test DISA_P05 |  |
| --- | --- |
| Image | SC image containing SMTPE pattern (height magnified by factor 2.0), Columns 512, Rows 1024, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\1024, Presentation Pixel Aspect Ratio 1\2 |
| Result | Display should be a normal SMPTE pattern (full screen) |

### Graphic Annotation

This subsection contains tests for all defined graphic types (ELLIPSE, INTERPOLATED, CIRCLE, POINT and POLYLINE).

Please note that the rendered output does not necessarily has to match exactly the associated result image in order to be standard compliant (e. g. it is not specified in the DICOM standard how to visualize a POINT or INTERPOLATED line).

| Test GRAN_P01 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: POLYLINE, Graphic Filled N, Graphic Data (128\256\192\128\320\128\ 384\256\320\384\192\384\128\256), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p02.res” |

| Test GRAN_P02 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: POLYLINE, Graphic Filled Y, Graphic Data (128\256\192\128\320\128\ 384\256\320\384\192\384\128\256), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p02.res” |

| Test GRAN_P03 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: POLYLINE, Graphic Filled N, Graphic Data (0.25\0.5\0.375\0.25\0.625\ 0.25\0.75\0.5\0.625\0.75\0.375\0.75\0.25\0.5), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p03.res” |

| Test GRAN_P04 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: POLYLINE, Graphic Filled Y, Graphic Data (0.25\0.5\0.375\0.25\0.625\ 0.25\0.75\0.5\0.625\0.75\0.375\0.75\0.25\0.5), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p04.res” |

| Test GRAN_P05 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: INTERPOLATED, Graphic Filled N, Graphic Data (128\256\192\128\ 320\128\384\256\320\384\192\384\128\256), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p05.res” (also see note above) |

| Test GRAN_P06 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: INTERPOLATED, Graphic Filled Y, Graphic Data (128\256\192\128\ 320\128\384\256\320\384\192\384\128\256), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p06.res” (also see note above) |

| Test GRAN_P07 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: INTERPOLATED, Graphic Filled N, Graphic Data (0.25\0.5\0.375\0.25\ 0.625\0.25\0.75\0.5\0.625\0.75\0.375\0.75\0.25\0.5), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p07.res” (also see note above) |

| Test GRAN_P08 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: INTERPOLATED, Graphic Filled Y, Graphic Data (0.25\0.5\0.375\0.25\ 0.625\0.25\0.75\0.5\0.625\0.75\0.375\0.75\0.25\0.5), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p08.res” (also see note above) |

| Test GRAN_P09 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: CIRCLE, Graphic Filled N, Graphic Data (256\256\384\256), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p09.res” |

| Test GRAN_P10 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: CIRCLE, Graphic Filled Y, Graphic Data (256\256\384\256), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p10.res” |

| Test GRAN_P11 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: CIRCLE, Graphic Filled N, Graphic Data (0.5\0.5\0.75\0.5), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p11.res” |

| Test GRAN_P12 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: CIRCLE, Graphic Filled N, Graphic Data (0.5\0.5\0.75\0.5), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p12.res” |

| Test GRAN_P13 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: ELLIPSE, Graphic Filled N, Graphic Data (128\256\384\256\256\192\ 256\320), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p13.res” |

| Test GRAN_P14 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: ELLIPSE, Graphic Filled Y, Graphic Data (128\256\384\256\256\192\ 256\320), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p14.res” |

| Test GRAN_P15 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: ELLIPSE, Graphic Filled N, Graphic Data (0.25\0.5\0.75\0.5\0.5\0.375\ 0.5\0.625), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p15.res” |

| Test GRAN_P16 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Type: ELLIPSE, Graphic Filled Y, Graphic Data (0.25\0.5\0.75\0.5\0.5\0.375\ 0.5\0.625), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p16.res” |

| Test GRAN_P17 |  |
| --- | --- |
| Image | SC image with five circles (burnt into the pixel data) |
| PState | Graphic Type: POINT, Graphic Data (128\256, 256\128, 256\256, 256\384, 384\256), Graphic Annotation Units PIXEL Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p17.res” |

| Test GRAN_P18 |  |
| --- | --- |
| Image | SC image with five circles (burnt into the pixel data) |
| PState | Graphic Type: POINT, Graphic Data (0.25\0.5, 0.5\0.25, 0.5\0.5, 0.5\0.75, 0.75\0.5), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p18.res” |

| Test GRAN_P19 |  |
| --- | --- |
| Image | SC black image |
| PState | Graphic Layer: LAYER1 (RecommendedDisplayGrayscaleValue 32767) - Graphic Type: CIRCLE, Graphic Filled Y, Graphic Data (0.5\0.5\0.6\0.5), Graphic Annotation Units DISPLAY Graphic Layer: LAYER2 (RecommendedDisplayGrayscaleValue 65535) - Graphic Type: CIRCLE, Graphic Filled N, Graphic Data (0.5\0.5\0.55\0.5), Graphic Annotation Units DISPLAY - Graphic Type: ELLIPSE, Graphic Filled N, Graphic Data (0.0\0.5\1.0\0.5\0.5\0.3\0.5\0.7), Graphic Annotation Units DISPLAY - Graphic Type: ELLIPSE, Graphic Filled N, Graphic Data (0.5\0.0\0.5\1.0\0.3\0.5\0.7\0.5), Graphic Annotation Units DISPLAY - Graphic Type: ELLIPSE, Graphic Filled N, Graphic Data (0.2\0.5\0.8\0.5\0.5\0.4\0.5\0.6), Graphic Annotation Units DISPLAY - Graphic Type: ELLIPSE, Graphic Filled N, Graphic Data (0.5\0.2\0.5\0.8\0.4\0.5\0.6\0.5), Graphic Annotation Units DISPLAY Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “gran_p19.res” |

### Text Annotation

This subsection contains tests for textual annotations.

Please note that the rendered output does not necessarily has to match exactly the associated result image in order to be standard compliant (e. g. it is not specified in the DICOM standard whether or not and how to show the Bounding Box of a textual annotation; some of the following test cases do show the Bounding Box and some do not).

| Test TEAN_P01 |  |
| --- | --- |
| Image | SC black image |
| PState | Bounding Box Annotation Units PIXEL, Unformatted Text Value “Text in bounding box of correct size”, Bounding Box Top Left Hand Corner 128\128, Bounding Box Bottom Right Hand Corner 320\144, Bounding Box Text Horizontal Justification LEFT Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p01.res” |

| Test TEAN_P02 |  |
| --- | --- |
| Image | SC black image |
| PState | Bounding Box Annotation Units DISPLAY, Unformatted Text Value “Text in bounding box of correct size”, Bounding Box Top Left Hand Corner 0.25\0.25, Bounding Box Bottom Right Hand Corner 0.625\0.28125, Bounding Box Text Horizontal Justification LEFT Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p02.res” |

| Test TEAN_P03 |  |
| --- | --- |
| Image | SC black image |
| PState | Bounding Box Annotation Units PIXEL, Unformatted Text Value “Too much text too fit in bounding box which isn't very large ...”, Bounding Box Top Left Hand Corner 128\128, Bounding Box Bottom Right Hand Corner 320\144, Bounding Box Text Horizontal Justification LEFT Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p03.res” |

| Test TEAN_P04 |  |
| --- | --- |
| Image | SC black image |
| PState | Bounding Box Annotation Units DISPLAY, Unformatted Text Value “Too much text too fit in bounding box which isn't very large ...”, Bounding Box Top Left Hand Corner 0.25\0.25, Bounding Box Bottom Right Hand Corner 0.625\0.28125 , Bounding Box Text Horizontal Justification LEFT Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p04.res” |

| Test TEAN_P05 |  |
| --- | --- |
| Image | SC black image |
| PState | Bounding Box Annotation Units PIXEL, Unformatted Text Value “Text in bounding box of correct size”, Bounding Box Top Left Hand Corner 128\128, Bounding Box Bottom Right Hand Corner 320\144, Bounding Box Text Horizontal Justification LEFT, Anchor Point Annotation Units PIXEL, Anchor Point 384\256, Anchor Point Visibility N Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p05.res” |

| Test TEAN_P06 |  |
| --- | --- |
| Image | SC black image |
| PState | Bounding Box Annotation Units DISPLAY, Unformatted Text Value “Text in bounding box of correct size”, Bounding Box Top Left Hand Corner 0.25\0.25, Bounding Box Bottom Right Hand Corner 0.625\0.28125, Bounding Box Text Horizontal Justification LEFT, Anchor Point Annotation Units DISPLAY, Anchor Point 0.75\0.5, Anchor Point Visibility N Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p06.res” |

| Test TEAN_P07 |  |
| --- | --- |
| Image | SC black image |
| PState | Bounding Box Annotation Units PIXEL, Unformatted Text Value “Text in bounding box of correct size”, Bounding Box Top Left Hand Corner 128\128, Bounding Box Bottom Right Hand Corner 320\144, Bounding Box Text Horizontal Justification LEFT, Anchor Point Annotation Units PIXEL, Anchor Point 384\256, Anchor Point Visibility Y Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_07.res” |

| Test TEAN_P08 |  |
| --- | --- |
| Image | SC black image |
| PState | Bounding Box Annotation Units DISPLAY, Unformatted Text Value “Text in bounding box of correct size”, Bounding Box Top Left Hand Corner 0.25\0.25, Bounding Box Bottom Right Hand Corner 0.625\0.28125, Bounding Box Text Horizontal Justification LEFT, Anchor Point Annotation Units DISPLAY, Anchor Point 0.75\0.5, Anchor Point Visibility Y Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p08.res” |

| Test TEAN_P09 |  |
| --- | --- |
| Image | SC black image |
| PState | Unformatted Text Value “Text with anchor point only”, Anchor Point Annotation Units PIXEL, Anchor Point 384\256, Anchor Point Visibility N Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p09.res” |

| Test TEAN_P10 |  |
| --- | --- |
| Image | SC black image |
| PState | Unformatted Text Value “Text with anchor point only”, Anchor Point Annotation Units DISPLAY, Anchor Point 0.75,0.5, Anchor Point Visibility N Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p10.res” |

| Test TEAN_P11 |  |
| --- | --- |
| Image | SC black image |
| PState | Unformatted Text Value “Text with anchor point only”, Anchor Point Annotation Units PIXEL, Anchor Point 384\256, Anchor Point Visibility Y Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p11.res” |

| Test TEAN_P12 |  |
| --- | --- |
| Image | SC black image |
| PState | Unformatted Text Value “Text with anchor point only”, Anchor Point Annotation Units DISPLAY, Anchor Point 0.75\0.5, Anchor Point Visibility Y Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p12.res” |

| Test TEAN_P13 |  |
| --- | --- |
| Image | SC black image |
| PState | Unformatted Text Value “Left justified, display relative ...” (five lines), Bounding Box Annotation Units DISPLAY, Bounding Box 0.0\0.0 0.5\0.5, Bounding Box Text Horizontal Justification LEFT, Anchor Point Annotation Units PIXEL, Anchor Point 256\256, Anchor Point Visibility Y Unformatted Text Value “Right justified, display relative ...” (five lines), Bounding Box Annotation Units DISPLAY, Bounding Box 0.5\0.0 1.0\0.5, Bounding Box Text Horizontal Justification RIGHT, Anchor Point Annotation Units PIXEL, Anchor Point 256\256, Anchor Point Visibility Y Unformatted Text Value “Centered, image relative text.”, Bounding Box Annotation Units PIXEL, Bounding Box 128\256 384\512, Bounding Box Text Horizontal Justification CENTER Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p13.res” |

| Test TEAN_P14 |  |
| --- | --- |
| Image | SC black image |
| PState | Unformatted Text Value “Normal text at the upper border.” (Bounding Box 1\1, 512\256), “Up-side down text at the lower border.” (512\512, 1\256), “Left rotated text at the left border.” (1\512, 256\1), “Right rotated text at the right border.” (512\1, 256\512); Anchor Point Annotation Units PIXEL, Bounding Box Text Horizontal Justification LEFT Presentation Size Mode: SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 |
| Result | Result on display should look like “tean_p14.res” |

### Display Shutter

This subsection contains tests for display shutters.  Please note that the shutters do not always exactly match the graphical forms in the images.  This is because of different algorithms for drawing circles, lines, etc. and deviations cannot be avoided.

| Test DISH_P01 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape CIRCULAR, Shutter Presentation Value 0 (black), Center of Circular Shutter 256\256, Radius of Circular Shutter 128 |
| Result | Result on display should look like “dish_p01.res” |

| Test DISH_P02 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape CIRCULAR, Shutter Presentation Value 65535 (white), Center of Circular Shutter 256\256, Radius of Circular Shutter 128 |
| Result | Result on display should look like “dish_p02.res” |

| Test DISH_P03 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape RECTANGULAR, Shutter Presentation Value 0 (black), Shutter Left Vertical Edge 128, Shutter Right Vertical Edge 384, Shutter Upper Horizontal Edge 128, Shutter Lower Horizontal Edge 384 |
| Result | Result on display should look like “dish_p03.res” |

| Test DISH_P04 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape RECTANGULAR, Shutter Presentation Value 65535 (white), Shutter Left Vertical Edge 128, Shutter Right Vertical Edge 384, Shutter Upper Horizontal Edge 128, Shutter Lower Horizontal Edge 384 |
| Result | Result on display should look like “dish_p04.res” |

| Test DISH_P05 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape POLYGONAL, Shutter Presentation Value 0 (black), Vertices of the Polygonal Shutter (256\128\128\192\128\320\256\384\384\320\384\192) |
| Result | Result on display should look like “dish_p05.res” |

| Test DISH_P06 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape POLYGONAL, Shutter Presentation Value 65535 (white), Vertices of the Polygonal Shutter (256\128\128\192\128\320\256\384\384\320\384\192) |
| Result | Result on display should look like “dish_p06.res” |

| Test DISH_P07 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape BITMAP, Shutter Presentation Value 0 (black), Overlay Columns 512, Overlay Rows 512, Overlay Type G, Overlay Origin 1\1, Overlay Bits Allocated 1, Overlay Position 0 |
| Result | Result on display should look like “dish_p07.res” |

| Test DISH_P08 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape BITMAP, Shutter Presentation Value 65535 (white), Overlay Columns 512, Overlay Rows 512, Overlay Type G, Overlay Origin 1\1, Overlay Bits Allocated 1, Overlay Position 0 |
| Result | Result on display should look like “dish_p08.res” |

| Test DISH_P09 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape POLYGONAL, Shutter Presentation Value 0 (black), Vertices of the Polygonal Shutter (257\133\233\199\169\169\199\233\133\257\199\281\169\345\233\ 315\257\381\281\315\345\345\315\281\381\257\315\233\345\169\281\199) |
| Result | Result on display should look like “dish_p09.res” |

| Test DISH_P10 |  |
| --- | --- |
| Image | SC image |
| PState | Shutter Shape POLYGONAL, Shutter Presentation Value 65535 (white), Vertices of the Polygonal Shutter (257\133\233\199\169\169\199\233\133\257\199\281\169\345\233\ 315\257\381\281\315\345\345\315\281\381\257\315\233\345\169\281\199) |
| Result | Result on display should look like “dish_p10.res” |

### Overlays

This subsection contains tests for overlays.

| Test OVLY_P01 |  |
| --- | --- |
| Image | SC image, Bits Allocated 16, Bits Stored 12, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs - Overlay (6000, XXXX): Overlay Columns 512, Overlay Rows 512, Overlay Type G, Overlay Origin 1\1, Overlay Bits Allocated 16, Overlay Position 15 - Overlay (6002, XXXX): Overlay Columns 512, Overlay Rows 512, Overlay Type G, Overlay Origin 1\1, Overlay Bits Allocated 16, Overlay Position 14 - Overlay (6004, XXXX): Overlay Columns 512, Overlay Rows 512, Overlay Type G, Overlay Origin 1\1, Overlay Bits Allocated 1, Overlay Position 0 - Overlay (6006, XXXX): Overlay Columns 512, Overlay Rows 512, Overlay Type G, Overlay Origin 1\1, Overlay Bits Allocated 1, Overlay Position 0 |
| PState | - Overlay (6000, XXXX): Overlay Activation Layer: LAYER1 - Overlay (6002, XXXX): Overlay Activation Layer: LAYER2 - Overlay (6004, XXXX): Overlay Activation Layer: LAYER3 - Overlay (6006, XXXX): Overlay Activation Layer: LAYER4 - Overlay (6008, XXXX): Overlay Columns 512, Overlay Rows 512, Overlay Type G, Overlay Origin 1\1, Overlay Bits Allocated 1, Overlay Position 0, Overlay Activation Layer: LAYER5 - Overlay (600a, XXXX): Overlay Columns 512, Overlay Rows 512, Overlay Type G, Overlay Origin 1\1, Overlay Bits Allocated 1, Overlay Position 0, Overlay Activation Layer: LAYER6 |
| Result | Result on display should look like “ovly_p01.res” |

### Complex Combination

This subsection contains more complex tests.

| Test CPLX_P01 |  |
| --- | --- |
| Image | SC image, Columns 1280, Rows 900, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to +255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, no LUTs |
| PState | - Displayed Area: Presentation Size Mode SCALE TO FIT, Displayed Area Top Left Hand Corner 768\388, Displayed Area Bottom Right Hand Corner 1280\900, Presentation Pixel Aspect Ratio 1\1 - Shutter: Shutter Shape CIRCULAR, Center Of Circular Shutter 644\1024, Radius of Cirular Shutter 128, Shutter Presentation Value 0 (black) - Spatial Transformations: Image Rotation 90, Image Horizontal Flip Y - Bounding Box: Annotation Units PIXEL, Unformatted Text Value “Image relative text overlapping shutter”, Bounding Box Top Left Hand Corner 896\516, Bounding Box Bottom Right Hand Corner 912\772, Bounding Box Text Horizontal Justification CENTER - Bounding Box: Annotation Units DISPLAY, Unformatted Text Value “Displayed Area relative text overlapping shutter”, Bounding Box Top Left Hand Corner 869\516, Bounding Box Bottom Right Hand Corner 0.75\0.78125, Bounding Box Text Horizontal Justification LEFT - Graphic Type: POLYLINE, Graphic Filled N, Graphic Data (960\452\960\836\976\836\976\452\960\452), Graphic Annotation Units PIXEL - Graphic Type: POLYLINE, Graphic Filled N, Graphic Data (0.125\0.59375\0.875\0.59375\0.875\0.625\0.125\0.625\0.125\0.59375), Graphic Annotation Units DISPLAY |
| Result | Result on display should look like “cplx_p01.res” |

| Test CPLX_P02 |  |
| --- | --- |
| Image | SC multi-frame image containing two “different” SMPTE patterns |
| PState | - Shutter: Shutter Shape RECTANGULAR, Shutter Presentation Value 0 (black), Shutter Left Vertical Edge 1, Shutter Right Vertical Edge 1024, Shutter Upper Horizontal Edge 32, Shutter Lower Horizontal Edge 512 Frame #1: - Displayed Area: Presentation Size Mode SCALE TO FIT, Displayed Area Top Left Hand Corner 513\1, Displayed Area Bottom Right Hand Corner 1024\512, Presentation Pixel Aspect Ratio 1\1 - Softcopy VOI LUT Sequence: Window Center = 50.5, Window Width = 51 - Bounding Box: Annotation Units PIXEL, Unformatted Text Value “Frame #1”, Bounding Box Top Left Hand Corner 513\1, Bounding Box Bottom Right Hand Corner 1024\128, Bounding Box Text Horizontal Justification LEFT Frame #2: - Displayed Area: Presentation Size Mode SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 1024\512, Presentation Pixel Aspect Ratio 2\1 - Bounding Box: Annotation Units PIXEL, Unformatted Text Value “Frame #2”, Bounding Box Top Left Hand Corner 1\1, Bounding Box Bottom Right Hand Corner 1024\128, Bounding Box Text Horizontal Justification LEFT |
| Result | Result on display should look like “cplx_p02.res” (2 frames) |

| Test CPLX_P03 |  |
| --- | --- |
| Image | SC image with SMPTE pattern, Bits Allocated 8, Bits Stored 8, Pixel Range 0 to 255, Pixel Representation is unsigned, Photometric Interpretation is MONOCHROME2, inappropriate VOI LUT transform (Window Center 0, Window Width 128), requiring identity VOI LUT transform |
| PState | Please note that the presentation state references two image objects. - Shutter: Shutter Shape RECTANGULAR, Shutter Presentation Value 0 (black), Shutter Left Vertical Edge 1, Shutter Right Vertical Edge 1024, Shutter Upper Horizontal Edge 32, Shutter Lower Horizontal Edge 512 Image #1 (“cplx_p03.dcm”): - Displayed Area: Presentation Size Mode SCALE TO FIT, Displayed Area Top Left Hand Corner 1\1, Displayed Area Bottom Right Hand Corner 512\512, Presentation Pixel Aspect Ratio 1\1 - VOI LUT transform not specified so assumed identity - Bounding Box: Annotation Units DISPLAY, Unformatted Text Value “Image #1”, Bounding Box Top Left Hand Corner 0.25\0.0, Bounding Box Bottom Right Hand Corner 0.75\0.25, Bounding Box Text Horizontal Justification CENTER Image #2 (“cplx_p02.dcm”): - Everything specified in “cplx_p02.pre” - Bounding Box: Annotation Units DISPLAY, Unformatted Text Value “Image #2”, Bounding Box Top Left Hand Corner 0.25\0.0, Bounding Box Bottom Right Hand Corner 0.75\0.25, Bounding Box Text Horizontal Justification CENTER |
| Result | Result on display should look like “cplx_p03.res” for image #1 and “cplx_p02.res” for image #2 (2 frames + centered text “Image #2” at the top of both frames) |

IHE Image Display Actors

Test Results

# Image Display Test Results

Application vendor

Name and software Version of the application

Person filling in this form, place, date

## Image Display without Presentation States

Please specify for each test whether the image has been rendered and displayed as described in chapter 3 (for some tests there is also a result file available).  If there are any comments please fill in the comment field.  It should be noted that all tests should be successful.

### Modality LUT

Please specify the test results for the Modality LUT transformation

| Success | Test Name | Comments |
| --- | --- | --- |
|  | MLUT_01 |  |
|  | MLUT_03 |  |
|  | MLUT_04 |  |
|  | MLUT_05 |  |
|  | MLUT_06 |  |
|  | MLUT_07 |  |
|  | MLUT_08 |  |
|  | MLUT_09 |  |
|  | MLUT_11 |  |
|  | MLUT_12 |  |
|  | MLUT_13 |  |
|  | MLUT_14 |  |
|  | MLUT_16 |  |
|  | MLUT_18 |  |
|  | MLUT_19 |  |

### VOI LUT

Please specify the test results for the VOI LUT transformation

| Success | Test Name | Comments |
| --- | --- | --- |
|  | VLUT_01 |  |
|  | VLUT_02 |  |
|  | VLUT_03 |  |
|  | VLUT_04 |  |
|  | VLUT_05 |  |
|  | VLUT_06 |  |
|  | VLUT_07 |  |
|  | VLUT_08 |  |
|  | VLUT_09 |  |
|  | VLUT_09 |  |
|  | VLUT_10 |  |
|  | VLUT_11 |  |

### Modality / VOI LUTs

Please specify the test results for the combination of Modality and VOI LUT.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | XLUT_01 |  |
|  | XLUT_02 |  |

## Image Display with Presentation States

Please specify for each test whether the Image and the applied Presentation State has been rendered and displayed as described in chapter 3.  If there are any comments please fill in the comment field.  It should be noted that there is no requirement that all tests are successful.

### Modality LUT

Please specify the test results for the Modality LUT transformation.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | MLUT_P01 |  |
|  | MLUT_P03 |  |
|  | MLUT_P04 |  |
|  | MLUT_P05 |  |
|  | MLUT_P06 |  |
|  | MLUT_P07 |  |
|  | MLUT_P08 |  |
|  | MLUT_P09 |  |
|  | MLUT_P11 |  |
|  | MLUT_P12 |  |
|  | MLUT_P13 |  |
|  | MLUT_P14 |  |
|  | MLUT_P16 |  |
|  | MLUT_P18 |  |
|  | MLUT_P19 |  |

### VOI LUT

Please specify the test results for the VOI LUT transformation.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | VLUT_P01 |  |
|  | VLUT_P02 |  |
|  | VLUT_P03 |  |
|  | VLUT_P04 |  |
|  | VLUT_P05 |  |
|  | VLUT_P06 |  |
|  | VLUT_P07 |  |
|  | VLUT_P08 |  |
|  | VLUT_P09 |  |
|  | VLUT_P10 |  |
|  | VLUT_P11 |  |
|  | VLUT_P12 |  |

### Presentation LUT

Please specify the test results for the Presentation LUT transformation.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | PLUT_P01 |  |
|  | PLUT_P02 |  |
|  | PLUT_P03 |  |
|  | PLUT_P04 |  |
|  | PLUT_P05 |  |
|  | PLUT_P06 |  |
|  | PLUT_P07 |  |
|  | PLUT_P08 |  |
|  | PLUT_P09 |  |
|  | PLUT_P10 |  |

### Modality / VOI / Presentation LUTs

Please specify the test results for the combination of different LUTs (Modality LUT, VOI LUT, Presentation LUT).

| Success | Test Name | Comments |
| --- | --- | --- |
|  | XLUT_P02 |  |
|  | XLUT_P03 |  |

### Spatial Transformations

Please specify the test results for the spatial transformations.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | SPAT_P01 |  |
|  | SPAT_P02 |  |
|  | SPAT_P03 |  |
|  | SPAT_P04 |  |
|  | SPAT_P05 |  |
|  | SPAT_P06 |  |
|  | SPAT_P07 |  |
|  | SPAT_P08 |  |
|  | SPAT_P09 |  |
|  | SPAT_P10 |  |
|  | SPAT_P11 |  |
|  | SPAT_P12 |  |
|  | SPAT_P13 |  |
|  | SPAT_P14 |  |
|  | SPAT_P15 |  |

### Displayed Area

Please specify the test results for the Displayed Area.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | DISA_P01 |  |
|  | DISA_P02 |  |
|  | DISA_P03 |  |
|  | DISA_P04 |  |
|  | DISA_P05 |  |

### Graphic Annotation

Please specify the test results for the graphic annotations.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | GRAN_P01 |  |
|  | GRAN_P02 |  |
|  | GRAN_P03 |  |
|  | GRAN_P04 |  |
|  | GRAN_P05 |  |
|  | GRAN_P06 |  |
|  | GRAN_P07 |  |
|  | GRAN_P08 |  |
|  | GRAN_P09 |  |
|  | GRAN_P10 |  |
|  | GRAN_P11 |  |
|  | GRAN_P12 |  |
|  | GRAN_P13 |  |
|  | GRAN_P14 |  |
|  | GRAN_P15 |  |
|  | GRAN_P16 |  |
|  | GRAN_P17 |  |
|  | GRAN_P18 |  |
|  | GRAN_P19 |  |

### Text Annotation

Please specify the test results for the text annotations.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | TEAN_P01 |  |
|  | TEAN_P02 |  |
|  | TEAN_P03 |  |
|  | TEAN_P04 |  |
|  | TEAN_P05 |  |
|  | TEAN_P06 |  |
|  | TEAN_P07 |  |
|  | TEAN_P08 |  |
|  | TEAN_P09 |  |
|  | TEAN_P10 |  |
|  | TEAN_P11 |  |
|  | TEAN_P12 |  |
|  | TEAN_P13 |  |
|  | TEAN_P14 |  |

### Display Shutter

Please specify the test results for the display shutters.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | DISH_P01 |  |
|  | DISH_P02 |  |
|  | DISH_P03 |  |
|  | DISH_P04 |  |
|  | DISH_P05 |  |
|  | DISH_P06 |  |
|  | DISH_P07 |  |
|  | DISH_P08 |  |
|  | DISH_P09 |  |
|  | DISH_P10 |  |

### Overlay

Please specify the test results for the overlays.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | OVLY_P01 |  |

### Complex Combination

Please specify the test results for the complex presentation states.

| Success | Test Name | Comments |
| --- | --- | --- |
|  | CPLX_P01 |  |
|  | CPLX_P02 |  |
|  | CPLX_P03 |  |

## General Comments

Please fill in general comments to the “Image Display Test”.

General Comments

# References

[1]	NEMA Standards Publications PS 3.14-2003, Digital Imaging and Communications in Medicine (DICOM), Part 14: Grayscale Standard Display Function, National Electrical Manufacturers Association, Rosslyn, VA, 2003.

[2]	NEMA Standards Publications PS 3.3-2003, Digital Imaging and Communications in Medicine (DICOM), Part 3: Information Object Definitions, National Electrical Manufacturers Association, Rosslyn, VA, 2003.