HIMMS and RSNA

IHE

Integrating the Healthcare Enterprise

Test Plan for Print Servers

2003-12-01

Marco Eichelberg1, Klaus Kleber2, Jörg Riesmeier1

1  Kuratorium OFFIS e.V., Escherweg 2, 26121 Oldenburg, Germany

2  VISUS Technology Transfer GmbH, Philippstraße 5, D-44803 Bochum, Germany

Contents

1	Introduction	3

2	Test Procedure for Print Servers	4

2.1	Prerequisites	4

2.2	Compiling the Print Server Test Cases from Source (optional)	4

2.3	Using Static Binaries on Posix Platforms (optional)	4

2.4	Configuring and Using the Print Server Test Cases	5

2.5	Configuring the Print Server Test Cases for TLS (IHE Basic Security)	6

2.5.1	The Test Patterns	7

2.6	The Test Cases for Transmissive Media (Film) Printers	8

2.6.1	Testing the GSDF Calibration	8

2.6.2	Testing the Visual Image Quality	9

2.6.3	Testing the General Print Protocol Implementation	10

2.6.4	Testing the Presentation LUT Implementation	12

2.6.5	Testing Printer Layouts	13

2.7	The Test Cases for Reflective Media (Paper) Printers	14

2.7.1	Testing the GSDF Calibration	14

2.7.2	Testing the Visual Image Quality	15

2.7.3	Testing the General Print Protocol Implementation	16

2.7.4	Testing the Presentation LUT Implementation	18

2.7.5	Testing Printer Layouts	20

2.8	Using the Excel Spreadsheet to Assess GSDF Conformance	21

2.8.1	Open Excel File and Fill Administrative Data	21

2.8.2	Measure Boxes and Enter Results	22

2.8.3	Print “Output” Page and Assess Deviation from GSDF	23

3	Print Server Test Results	25

3.1	Testing the GSDF Calibration	25

3.2	Testing the Visual Image Quality	26

3.3	Testing the General Print Protocol Implementation	26

3.4	Testing the Presentation LUT Implementation	26

3.5	Testing the Presentation LUT Implementation	27

3.6	General Comments	27

4	References	27

# Introduction

This document is intended for those vendors participating in IHE which supply Print Server actors for the “Consistent Presentation of Images” Integration Profile.  Print Server Actors are required to implement the IHE Transaction 23 “Print Request with Presentation LUT” of the IHE Radiology Technical Framework. This document specifies a number of test cases which allow to assess

- whether the Print Server is correctly calibrated to the DICOM Grayscale Standard Display Function (GSDF) [1] as required by the Technical Framework,

- whether the implementation of the Presentation LUT SOP Class leads to the expected results

- whether the use of Presentation LUTs leads to visible artifacts in the hardcopies that affect image quality

- how the Print Server reacts on different optional elements of the Print Management Service Class.  Successful performance of these tests is not a prerequisite for successful completion of this test plan, but significantly increases the probability that the Print Server will interoperate with different kinds of Print Composer Actors.

Section 2 describes prerequisites of this test plan, how to set up and perform the test cases, and how to use the Excel spreadsheet that is used to assess the measured results from many of the test cases.  Section 3 is a form to be filled in with a Print Server’s test results.

The tools described in this document may also be used with Print Server actors supporting the IHE “Basic Security” Integration Profile, i. e. communicating DICOM over TLS.

If a vendor supplies more than one Print Server, or if the Print Server can be operated in different modes which exhibit different behavior on the DICOM network protocol level, the tests should be run for each Print Server or mode of operation, and multiple copies of the form should be filled in.  Completed forms are for your convenience and do not need to be turned in.

# Test Procedure for Print Servers

## Prerequisites

The following items are required for performance of the Test Plan:

- Densitometer for measuring the optical density of a printed film (transmissive media) or paper (reflective media).  Suitable devices for this purpose are identified in a separate document entitled “Calibration Tools”.

- Print Server Test Cases consisting of a number of print jobs that will be printed during the execution of the test plan.  These are provided as part of the IHE MESA tools distribution.

- Microsoft Excel 97 Spreadsheet “gsdf_od.xls”. This spreadsheet is explained in section 2.4.  Also provided as part of the IHE MESA tools distribution.

## Compiling the Print Server Test Cases from Source (optional)

The Print Server Test Cases are provided in pre-compiled form for the most popular platforms.  If you prefer to compile from source, of if you need to use the Print Server Test Cases on a platform for which no pre-compiled binaries are available, proceed as follows: Download and unpack a binary distribution of the Print Server Test Cases. When unpacked, the distribution creates a directory (which may be a subdirectory of the total MESA distribution) called printscp_testcases.  This directory contains the test patterns and configuration files required to perform the Print Server test cases.  It also contains two applications (binaries) in the bin subdirectory:

dbregimg  (dbregimg.exe on Windows)

dcmprscu  (dcmprscu.exe on Windows)

These binaries have to be replaced.  The source code is provided with the OFFIS DICOM Toolkit DCMTK, which is available for download from http://dicom.offis.de/.  Follow the compile instructions provided with the toolkit and make sure that you compile with Zlib support and OpenSSL support enabled.  After compilation, the dbregimg binary can be found in dcmtk/imagectn/tests and the dcmprscu binary can be found in dcmtk/dcmpstat/apps.  If the DCMTK tools have been installed with “make install” according to the installation instructions, both tools should be present in the installation directory, e. g. /usr/local/dicom/bin.  In addition to these two binaries, the file dicom.dic is required in the lib subdirectory on Unix platforms.  This file is also contained in the DCMTK DICOM toolkit in dcmtk/dcmdata/libsrc and is installed during “make install” in /usr/local/dicom/lib.

## Using Static Binaries on Posix Platforms (optional)

The pre-compiled Print Server Test Cases for Posix platforms (e. g. Linux) contain two sets of binaries, one dynamically linked and one statically linked.  The dynamically linked binaries are used by default.  If you cannot run these binaries due to missing shared objects (system libraries), it may be worthwhile to try the statically linked binaries. In the bin/ directory of the distribution, just rename the binaries as follows:

mv dcmprscu_static dcmprscu

mv dbregimg_static dbregimg

If these binaries also do not work, you need to re-compile from source as explained in section 2.2 above.

## Configuring and Using the Print Server Test Cases

Load the configuration file printscp_test.cfg that is distributed as part of the Print Server test cases into your text editor and to adapt it for your local Print Server settings.  You will the following comment in the configuration file five times, with different numbers for X:

# --- adjust the next X lines to match your print SCPs characteristics

Please adapt these settings according to the network address and capabilities of the Print Server under test (see table below), and leave everything else unchanged.  Save the updated configuration file.

| Keyword | Value | Description |
| --- | --- | --- |
| Hostname | string | The IP number or domain name of the Print Server in conventional TCP/IP notation. |
| Port | 1..65535 | IP port number under which the Print SCP receives DICOM associations. |
| AETitle | string | DICOM application entity title used by the Print SCP to identify itself during association negotiation. |
| Supports12Bit | true or false | Value should be “true” if the Print Server accepts image pixel data with 12 bits/pixel. If the Print Server only accepts image pixel data with 8 bits/pixel, the value should be “false”. Please note that this setting is only related to the Print Server’s ability to handle 12 bit image data on the DICOM network protocol level. It is not related to the ability of the printer hardware to process 12 bit image data. |
| PresentationLUTinFilmSession | true or false | The 1999 edition of the DICOM standard contains an inconsistency in the definition of the Presentation LUT SOP class. The attributes "Referenced Presentation LUT", "Illumination" and "Reflected Ambient Light" can either be part of the Basic Film Session or be part of the Basic Film Box. DICOM Correction Proposal CP 173 defines that these attributes have to be specified on Basic Film Box level, and this is also a requirement according to the Technical Framework. Nevertheless, if the Print Server only accepts these elements on Film Session level, this behavior can be activated in the Print SCU by setting this value to “true”. |

The test cases are shipped with a Shell script (for Unix systems) and a Batch file (for Windows systems) which allow to perform individual test cases or all of the tests in a single batch.  The script is invoked as “printscp_test”, followed by a single parameter:

- printscp_test reindex causes the operating system dependent index file required by the Print SCU to be created.  This should be performed once before the test cases are sent to the printer.

- printscp_test 001 performs test case #001.  Numbers must be specified with exactly three digits.  All test case numbers defined in this document can be used.

- printscp_test all_film performs all test cases for transmissive media (film) printers.

- printscp_test all_paper performs all test cases for reflective media (paper) printers.

For each performed test case a log file is created in the log subdirectory.  The filename is XXX.log, where XXX is the number of the test case in three digit notation.  This is a plain text file that can be viewed and printed with any text editor.  It contains a detailed log of all DICOM network communication performed between the Print SCU and the Print Server under test.  This file should be consulted if any problems occur during performance of the test cases, e. g. if no hardcopies are printed.

## Configuring the Print Server Test Cases for TLS (IHE Basic Security)

The Print Server Test Cases may be used to test Print Server actors that support IHE Basic Security Integration Profile as a Secure Node actor.  For this purpose, TLS support has to be enabled in the configuration file printscp_test.cfg. You will the following comment in the configuration file five times:

# --- adjust the next line if you want to enable TLS (IHE Secure Node)

In all cases, the next line is “UseTLS = false”.  It should be changed to “UseTLS = true” which enables TLS communication.

The Print Server Test Cases will negotiate two alternative TLS cipher suites: TLS_RSA_WITH_3DES_EDE_CBC_SHA (the DICOM default ciphersuite, using 168 bit 3DES encryption)  and TLS_RSA_WITH_NULL_SHA (unencrypted, support required by IHE).  Encrypted transmission is preferred in the ciphersuite negotiation.

The Print Server Test Cases do not to perform any validation of the Print Server’s certificate.  It might be necessary, however, to change the certificate and private key the that Test Cases present to the Print Server actor in order to authenticate themselves.  The certificate is contained in file cert.pem in the tls/ directory, the private key (which has to be unencrypted) is contained in file pkey.pem in the same directory.  Both files are encoded in PEM format.  The certificate provided with this release will expire in September, 2006.  The certificate and key files can be replaced by versions provided by the user.

### The Test Patterns

The Print Server test cases make use of four test patterns which are printed with different parameters.

| DSITest Pattern | Wedge Pattern |
| --- | --- |
| 12-Bit Wedge Pattern | Pixel Aspect Ratio Pattern |

- The DSITest Pattern contains 16 boxes with equidistant shades of gray from black to white.  In addition it shows 2% contrast boxes.

- The Wedge Pattern contains a transition from black to white in 45 degree angled lines.  This pattern only contains 256 different shades of gray.

- The 12-Bit Wedge Pattern contains four stripes with a transition from black to white in 45 degree angled lines.  This pattern contains 4096 shades of gray.

- The Pixel Aspect Ratio pattern shows a transition from black to white.  The source image contains a pixel aspect ratio of 4:1 (512 by 128 pixels).  Correctly printed, the image has a square shape.

## The Test Cases for Transmissive Media (Film) Printers

The tests specified in this section are intended for transmissive media (film) printers only.  An identical set of test for reflective media (paper) printers is provided in the following section.  The test sets only differ in the use of the “Illumination” and “Reflected Ambient Light” attributes and in the maximum and minimum optical density, which differ for film and paper.

### Testing the GSDF Calibration

The first set of tests examines whether the Print Server correctly creates GSDF calibrated hardcopies.  Each test sends a single image, a Presentation LUT and well-defined lighting conditions to the Print Server.

Test #001: GSDF Curve under Default Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are typical for screen/film systems.

Evaluation

The hardcopy should match the GSDF for Illumination=2000, Reflected Ambient Light=10. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #002: GSDF Curve under unusual Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are unusual for screen/film systems – the illumination is rather dim and the reflection caused by ambient light is very high. The Print Server must be able to match the GSDF for these values.

Evaluation

The hardcopy should match the GSDF for Illumination=1000, Reflected Ambient Light=100. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #003: GSDF Curve with Presentation LUT (256 Entries), Default Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are typical for screen/film systems. This test uses a real Presentation LUT (look-up table) with 256 entries instead of a Presentation LUT Shape. It is sufficient if either test #003 or test #004 is passed.

Evaluation

The hardcopy should match the GSDF for Illumination=2000, Reflected Ambient Light=10. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #004: GSDF Curve with Presentation LUT (4096 Entries), Default Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are typical for screen/film systems. This test uses a real Presentation LUT (look-up table) with 4096 entries instead of a Presentation LUT Shape. It is sufficient if either test #003 or test #004 is passed.

Evaluation

The hardcopy should match the GSDF for Illumination=2000, Reflected Ambient Light=10. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #005: GSDF Curve with Presentation LUT (256 Entries), unusual Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are unusual for screen/film systems – the illumination is rather dim and the reflection caused by ambient light is very high. The Print Server must be able to match the GSDF for these values. This test uses a real Presentation LUT (look-up table) with 256 entries instead of a Presentation LUT Shape. It is sufficient if either test #005 or test #006 is passed.

Evaluation

The hardcopy should match the GSDF for Illumination=1000, Reflected Ambient Light=100. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #006: GSDF Curve with Presentation LUT (4096 Entries), unusual Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are unusual for screen/film systems – the illumination is rather dim and the reflection caused by ambient light is very high. The Print Server must be able to match the GSDF for these values. This test uses a real Presentation LUT (look-up table) with 4096 entries instead of a Presentation LUT Shape. It is sufficient if either test #005 or test #006 is passed.

Evaluation

The hardcopy should match the GSDF for Illumination=1000, Reflected Ambient Light=100. Use Excel spreadsheet described in section 2.4 to assess conformance.

### Testing the Visual Image Quality

The second set of tests examines whether the Print SCP’s Presentation LUT implementation affects the image quality on the rendered film.  Each test sends an image and a Presentation LUT to the printer.  The resulting film is checked for aliasing artifacts (“edges” or “diagonal lines”) visible on the film, which indicate that the display curve of the Print SCP is not smooth enough to produce a high quality result.

Test #010: Test for Visual Artifacts, 8-bit image data with Presentation LUT Shape

This test print contains the wedge pattern (with 256 different shades of gray) printed under default lighting conditions with an IDENTITY Presentation LUT Shape.

Evaluation

This test is evaluated visibly (by a human observer). The hardcopy should not contain visible contours – the transition from dark to bright within the wedge should be smooth. Diagonal lines visible on the wedge indicate a reduced image quality resulting either from a loss of shades of gray as a result of the image processing in the printer, or from “steps” in the display curve.

Test #011: Test for Visual Artifacts, 8-bit image data with Presentation LUT (256 entries)

This test print contains the wedge pattern (with 256 different shades of gray) printed under default lighting conditions. This test uses a real Presentation LUT (look-up table) with 256 entries instead of a Presentation LUT Shape.

Evaluation

This test is evaluated visibly (by a human observer). The hardcopy should not contain visible contours – the transition from dark to bright within the wedge should be smooth. Diagonal lines visible on the wedge indicate a reduced image quality resulting either from a loss of shades of gray as a result of the image processing in the printer, or from “steps” in the display curve. The Presentation LUT used for this test should result in a hardcopy that is slightly brighter in the mid-tone areas than the result of test #10.

Test #012 has been retired.

Test #013: Test for Visual Artifacts, 12-bit image data with Presentation LUT Shape

This test print contains the 12-bit wedge pattern (with 4096 different shades of gray) printed under default lighting conditions with an IDENTITY Presentation LUT Shape.

Evaluation

This test is evaluated visibly (by a human observer). The hardcopy should not contain visible contours – the transition from dark to bright within the wedge should be smooth. Diagonal lines visible on the wedge indicate a reduced image quality resulting either from a loss of shades of gray as a result of the image processing in the printer, or from “steps” in the display curve.

Test #014 has been retired.

Test #015: Test for Visual Artifacts, 12-bit image data with Presentation LUT (4096 entries)

This test print contains the 12-bit wedge pattern (with 4096 different shades of gray) printed under default lighting conditions. This test uses a real Presentation LUT (look-up table) with 4096 entries instead of a Presentation LUT Shape.

Evaluation

This test is evaluated visibly (by a human observer). The hardcopy should not contain visible contours – the transition from dark to bright within the wedge should be smooth. Diagonal lines visible on the wedge indicate a reduced image quality resulting either from a loss of shades of gray as a result of the image processing in the printer, or from “steps” in the display curve. The Presentation LUT used for this test should result in a hardcopy that is slightly brighter in the mid-tone areas than the result of test #10.

### Testing the General Print Protocol Implementation

The third set of tests sends a couple of print jobs to the Print SCP, each of which exercises different aspects of the DICOM print protocol.

Test #020: Referenced Presentation LUT SQ on Film Session Level

This test print sends a Referenced Presentation LUT Sequence as part of the Basic Film Session N-CREATE request. The interesting question is, whether the Print Server accepts this element and whether a Referenced Presentation LUT Sequence is sent back as part of the N-CREATE response message created by the Print Server.

Evaluation

The log file for this test case (log/020.log) should be examined. If the Print Server accepts the Referenced Presentation LUT Sequence and sends back a Referenced Presentation LUT Sequence as part of the Basic Film Session N-CREATE request, the Print Server implementation does not implement DICOM Correction Proposal 173. If the N-CREATE response contains an error status code, or if the Basic Film Session N-CREATE response does not contain a Referenced Presentation LUT Sequence but the Basic Film Box N-CREATE response does, the Print Server implements DICOM Correction Proposal 173. The test is passed if the Print Server supports CP 173. Note: There might be Print SCP implementations which fully support Presentation LUT both on Basic Film Box and Basic Film Session level (i. e. both CP 173 and the pre-CP 173 behavior). On such implementations, Test #020 will fail and Test #021 will succeed. For this reason, this test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #021: Referenced Presentation LUT SQ on Film Box Level

This test print sends a Referenced Presentation LUT Sequence as part of the Basic Film Box N-CREATE request. The interesting question is, whether the Print Server accepts this element and whether a Referenced Presentation LUT Sequence is sent back as part of the N-CREATE response message created by the Print Server.

Evaluation

The log file for this test case (log/020.log) should be examined. If the Print Server accepts the Referenced Presentation LUT Sequence and sends back a Referenced Presentation LUT Sequence as part of the Basic Film Box N-CREATE request, the Print Server implementation supports DICOM Correction Proposal 173. If the N-CREATE response contains an error status code, or if the Basic Film Box N-CREATE response does not contain a Referenced Presentation LUT Sequence but the Basic Film Session N-CREATE response does, the Print Server does not implement DICOM Correction Proposal 173. The test is passed if the Print Server supports CP 173.

Test #022: Landscape Orientation

This test prints the dsitest pattern in landscape orientation.

Evaluation

The hardcopy should display the dsitest pattern in landscape orientation. If no hardcopy is created, or if the dsitest pattern is printed in portrait orientation, the Print Server does not support landscape orientation and the test is not passed.

Test #023: Maximum Density

This test print contains the dsitest pattern printed under lighting conditions which are typical for screen/film systems. However, the Print Server is requested to print the image with a maximum density of 1.5 OD instead of the printer default (which is around 3.0 OD on many printers).

Evaluation

The hardcopy should match the GSDF for Illumination=2000, Reflected Ambient Light=10. The maximum density should be 1.5 OD. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #024: Minimum Density (optional)

This test print contains the dsitest pattern printed under lighting conditions which are typical for screen/film systems. However, the Print Server is requested to print the image with a minimum density of 1.0 OD instead of the printer default (which is around 0.2 OD on many printers).

Evaluation

The hardcopy should match the GSDF for Illumination=2000, Reflected Ambient Light=10. The minimum density should be 1.0 OD. Use Excel spreadsheet described in section 2.4 to assess conformance. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #025: Collation (optional)

This test prints the dsitest pattern using an N-ACTION request on Basic Film Session level (instead of the Basic Film Box level which is used in all other tests).

Evaluation

The result should equal the result from test #001. If no hardcopy is printed, the Print Server does not support collation and the test is not passed. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #026: 8-bit Image Transmission (optional)

This test prints the dsitest pattern. The pattern is transmitted with 8 bits/pixel in the Basic Grayscale Image Box.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #027: 12-bit Image Transmission (optional)

This test prints the dsitest pattern. The pattern is transmitted with 12 bits/pixel in the Basic Grayscale Image Box.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #028: Pixel Aspect Ratio

This test prints the Pixel Aspect Ratio Test Pattern with a pixel aspect ratio of 4:1.

Evaluation

The hardcopy should display a square image. If the image is a rectangle, the Print Server does not support Pixel Aspect Ratio and the test is not passed.

Test #029: Requested Image Size (optional)

This test prints the dsitest pattern with a requested image size of 127mm (5”).

Evaluation

The hardcopy should display the dsitest image with a size of 127x127mm (5x5”). If no hardcopy is created, or if the image is printed with a size different from 127x127mm, the Print Server does not support Requested Image Size and the test is not passed. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #030: Polarity (optional)

This test sends an inverted dsitest pattern to the printer and sets polarity to “REVERSE”.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #031: MONOCHROME 1 (optional)

This test sends a dsitest pattern using MONOCHROME1 photometric interpretation to the printer.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

### Testing the Presentation LUT Implementation

This set of tests checks details of the Print SCPs implementation of the Presentation LUT SOP Class.

Test #040: LIN OD Presentation LUT Shape (optional)

This test sends a dsitest pattern that has been converted to a value space linear to optical density. A Presentation LUT Shape of “LIN OD” is used. The image has been processed to be linear to the optical densities of a hardcopy printed with Dmax = 3.0 OD, Dmin = 0.2 OD, on a GSDF calibrated printer with Illumination = 2000 cd/m2 and Reflected Ambient Light = 10 cd/m2.

Evaluation

The result should equal the result from test #001, or at least be very similar. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #041: Presentation LUT with 256 entries, Image with 8 bits/pixel (optional)

This test print creates a Presentation LUT with 256 entries and transmits the dsitest pattern with 8 bits/pixel.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #042: Presentation LUT with 4096 entries, Image with 12 bits/pixel (optional)

This test print creates a Presentation LUT with 4096 entries and transmits the dsitest pattern with 12 bits/pixel.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #043: Presentation LUT with 256 entries, Image with 12 bits/pixel (optional)

This test print creates a Presentation LUT with 256 entries and transmits the dsitest pattern with 12 bits/pixel. Combinations of Presentation LUT and pixel data where the number of LUT entries does not match with the possible number of grayscale values in the pixel data have been declared a protocol violation in CP 202 (2001) but were allowed in earlier versions of the DICOM standard. Therefore, Print Servers may reject this test.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #044: Presentation LUT with 4096 entries, Image with 8 bits/pixel (optional)

This test print creates a Presentation LUT with 4096 entries and transmits the dsitest pattern with 8 bits/pixel. Combinations of Presentation LUT and pixel data where the number of LUT entries does not match with the possible number of grayscale values in the pixel data have been declared a protocol violation in CP 202 (2001) but were allowed in earlier versions of the DICOM standard. Therefore, Print Servers may reject this test.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #045: Presentation LUT with 256 entries and 11 bits/entry (optional)

This test print creates a Presentation LUT with 256 entries and transmits the dsitest pattern with 8 bits/pixel. The Presentation LUT is unusual because the number of significant bits per entry is 11.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #046: Presentation LUT with 256 entries and 13 bits/entry (optional)

This test print creates a Presentation LUT with 256 entries and transmits the dsitest pattern with 8 bits/pixel. The Presentation LUT is unusual because the number of significant bits per entry is 13.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #047: Scrambled Presentation LUT and Image (optional)

This test print creates a Presentation LUT with 4096 entries and transmits the dsitest pattern with 12 bits/pixel. Both Presentation LUT and Image pixel data are scrambled such that only a correct application of the Presentation LUT restores a visible image.

Evaluation

The result should equal the result from test #001. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #048: Presentation LUT with OW value representation (optional)

This test print creates a Presentation LUT with 4096 entries and transmits the dsitest pattern with 12 bits/pixel. The Presentation LUT is unusual because the DICOM Value Representation of the LUT is “OW”. This value representation has been introduced for LUT Data in DICOM Supplement 33 in 1999.

Evaluation

The log file for this test case (log/048.log) should be examined. If the Print Server accepts the Presentation LUT N-CREATE request, it may be assumed that the Print Server supports LUT Data with OW value representation. Otherwise, the test is not passed. This test is optional. Passing the test is not a requirement for IHE Print Servers.

### Testing Printer Layouts

This set of tests tries out the most common image display formats (page layouts). All tests in this section are optional.

Test #050: 2x2 Layout (optional)

This test creates and prints a Basic Film Box with an image display format (page layout) of “STANDARD\2,2”.

Evaluation

The result should show the dsitest and the wedge pattern tiled on a print-out with 2x2 image boxes. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #051: 3x5 Layout (optional)

This test creates and prints a Basic Film Box with an image display format (page layout) of “STANDARD\2,2”.

Evaluation

The result should show the dsitest and the wedge pattern tiled on a print-out with 3x5 image boxes. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #052: 4x5 Layout (optional)

This test creates and prints a Basic Film Box with an image display format (page layout) of “STANDARD\2,2”.

Evaluation

The result should show the dsitest and the wedge pattern tiled on a print-out with 4x5 image boxes. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #053: 4x6 Layout (optional)

This test creates and prints a Basic Film Box with an image display format (page layout) of “STANDARD\2,2”.

Evaluation

The result should show the dsitest and the wedge pattern tiled on a print-out with 4x6 image boxes. This test is optional. Passing the test is not a requirement for IHE Print Servers.

## The Test Cases for Reflective Media (Paper) Printers

The tests specified in this section are intended for reflective media (paper) printers only.  An identical set of test for transmissive media (film) printers is provided in the previous section. The test sets only differ in the use of the “Illumination” and “Reflected Ambient Light” attributes and in the maximum and minimum optical density, which differ for film and paper.

### Testing the GSDF Calibration

The first set of tests examines whether the Print Server correctly creates GSDF calibrated hardcopies.  Each test sends a single image, a Presentation LUT and well-defined lighting conditions to the Print Server.

Test #101: GSDF Curve under Default Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are typical for reflective media.

Evaluation

The hardcopy should match the GSDF for Illumination=150, Reflected Ambient Light=0. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #102: GSDF Curve under unusual Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under unusual lighting conditions for reflective media – the illumination is extremely bright. The Print Server must be able to match the GSDF for these values.

Evaluation

The hardcopy should match the GSDF for Illumination=500, Reflected Ambient Light=0. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #103: GSDF Curve with Presentation LUT (256 Entries), Default Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are typical for reflective media. This test uses a real Presentation LUT (look-up table) with 256 entries instead of a Presentation LUT Shape. It is sufficient if either test #103 or test #104 is passed.

Evaluation

The hardcopy should match the GSDF for Illumination=150, Reflected Ambient Light=0. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #104: GSDF Curve with Presentation LUT (4096 Entries), Default Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are typical for reflective media. This test uses a real Presentation LUT (look-up table) with 4096 entries instead of a Presentation LUT Shape. It is sufficient if either test #103 or test #104 is passed.

Evaluation

The hardcopy should match the GSDF for Illumination=150, Reflected Ambient Light=0. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #105: GSDF Curve with Presentation LUT (256 Entries), unusual Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are unusual for reflective media – the illumination is extremely bright. The Print Server must be able to match the GSDF for these values. This test uses a real Presentation LUT (look-up table) with 256 entries instead of a Presentation LUT Shape. It is sufficient if either test #105 or test #106 is passed.

Evaluation

The hardcopy should match the GSDF for Illumination=500, Reflected Ambient Light=0. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #106: GSDF Curve with Presentation LUT (4096 Entries), unusual Lighting Conditions

This test print contains the dsitest pattern (see section 2.4) printed under lighting conditions which are unusual for reflective media – the illumination is extremely bright. The Print Server must be able to match the GSDF for these values. This test uses a real Presentation LUT (look-up table) with 4096 entries instead of a Presentation LUT Shape. It is sufficient if either test #105 or test #106 is passed.

Evaluation

The hardcopy should match the GSDF for Illumination=500, Reflected Ambient Light=0. Use Excel spreadsheet described in section 2.4 to assess conformance.

### Testing the Visual Image Quality

The second set of tests examines whether the Print SCP’s Presentation LUT implementation affects the image quality on the rendered film.  Each test sends an image and a Presentation LUT to the printer.  The resulting film is checked for aliasing artifacts (“edges” or “diagonal lines”) visible on the film, which indicate that the display curve of the Print SCP is not smooth enough to produce a high quality result.

Test #110: Test for Visual Artifacts, 8-bit image data with Presentation LUT Shape

This test print contains the wedge pattern (with 256 different shades of gray) printed under default lighting conditions with an IDENTITY Presentation LUT Shape.

Evaluation

This test is evaluated visibly (by a human observer). The hardcopy should not contain visible contours – the transition from dark to bright within the wedge should be smooth. Diagonal lines visible on the wedge indicate a reduced image quality resulting either from a loss of shades of gray as a result of the image processing in the printer, or from “steps” in the display curve.

Test #111: Test for Visual Artifacts, 8-bit image data with Presentation LUT (256 entries)

This test print contains the wedge pattern (with 256 different shades of gray) printed under default lighting conditions. This test uses a real Presentation LUT (look-up table) with 256 entries instead of a Presentation LUT Shape.

Evaluation

This test is evaluated visibly (by a human observer). The hardcopy should not contain visible contours – the transition from dark to bright within the wedge should be smooth. Diagonal lines visible on the wedge indicate a reduced image quality resulting either from a loss of shades of gray as a result of the image processing in the printer, or from “steps” in the display curve. The Presentation LUT used for this test should result in a hardcopy that is slightly brighter in the mid-tone areas than the result of test #110.

Test #112 has been retired.

Test #113: Test for Visual Artifacts, 12-bit image data with Presentation LUT Shape

This test print contains the 12-bit wedge pattern (with 4096 different shades of gray) printed under default lighting conditions with an IDENTITY Presentation LUT Shape.

Evaluation

This test is evaluated visibly (by a human observer). The hardcopy should not contain visible contours – the transition from dark to bright within the wedge should be smooth. Diagonal lines visible on the wedge indicate a reduced image quality resulting either from a loss of shades of gray as a result of the image processing in the printer, or from “steps” in the display curve.

Test #114 has been retired.

Test #115: Test for Visual Artifacts, 12-bit image data with Presentation LUT (4096 entries)

This test print contains the 12-bit wedge pattern (with 4096 different shades of gray) printed under default lighting conditions. This test uses a real Presentation LUT (look-up table) with 4096 entries instead of a Presentation LUT Shape.

Evaluation

This test is evaluated visibly (by a human observer). The hardcopy should not contain visible contours – the transition from dark to bright within the wedge should be smooth. Diagonal lines visible on the wedge indicate a reduced image quality resulting either from a loss of shades of gray as a result of the image processing in the printer, or from “steps” in the display curve. The Presentation LUT used for this test should result in a hardcopy that is slightly brighter in the mid-tone areas than the result of test #110.

### Testing the General Print Protocol Implementation

The third set of tests sends a couple of print jobs to the Print SCP, each of which exercises different aspects of the DICOM print protocol.

Test #120: Referenced Presentation LUT SQ on Film Session Level

This test print sends a Referenced Presentation LUT Sequence as part of the Basic Film Session N-CREATE request. The interesting question is, whether the Print Server accepts this element and whether a Referenced Presentation LUT Sequence is sent back as part of the N-CREATE response message created by the Print Server.

Evaluation

The log file for this test case (log/020.log) should be examined. If the Print Server accepts the Referenced Presentation LUT Sequence and sends back a Referenced Presentation LUT Sequence as part of the Basic Film Session N-CREATE request, the Print Server implementation does not implement DICOM Correction Proposal 173. If the N-CREATE response contains an error status code, or if the Basic Film Session N-CREATE response does not contain a Referenced Presentation LUT Sequence but the Basic Film Box N-CREATE response does, the Print Server implements DICOM Correction Proposal 173. The test is passed if the Print Server supports CP 173. Note: There might be Print SCP implementations which fully support Presentation LUT both on Basic Film Box and Basic Film Session level (i. e. both CP 173 and the pre-CP 173 behavior). On such implementations, Test #120 will fail and Test #121 will succeed. For this reason, this test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #121: Referenced Presentation LUT SQ on Film Box Level

This test print sends a Referenced Presentation LUT Sequence as part of the Basic Film Box N-CREATE request. The interesting question is, whether the Print Server accepts this element and whether a Referenced Presentation LUT Sequence is sent back as part of the N-CREATE response message created by the Print Server.

Evaluation

The log file for this test case (log/020.log) should be examined. If the Print Server accepts the Referenced Presentation LUT Sequence and sends back a Referenced Presentation LUT Sequence as part of the Basic Film Box N-CREATE request, the Print Server implementation supports DICOM Correction Proposal 173. If the N-CREATE response contains an error status code, or if the Basic Film Box N-CREATE response does not contain a Referenced Presentation LUT Sequence but the Basic Film Session N-CREATE response does, the Print Server does not implement DICOM Correction Proposal 173. The test is passed if the Print Server supports CP 173.

Test #122: Landscape Orientation

This test prints the dsitest pattern in landscape orientation.

Evaluation

The hardcopy should display the dsitest pattern in landscape orientation. If no hardcopy is created, or if the dsitest pattern is printed in portrait orientation, the Print Server does not support landscape orientation and the test is not passed.

Test #123: Maximum Density

This test print contains the dsitest pattern printed under lighting conditions which are typical for reflective media. However, the Print Server is requested to print the image with a maximum density of 0.7 OD instead of the printer default (which is around 1.5 OD on typical paper printers).

Evaluation

The hardcopy should match the GSDF for Illumination=150, Reflected Ambient Light=0. The maximum density should be 0.7 OD. Use Excel spreadsheet described in section 2.4 to assess conformance.

Test #124: Minimum Density (optional)

This test print contains the dsitest pattern printed under lighting conditions which are typical for reflective media. However, the Print Server is requested to print the image with a minimum density of 0.4 OD instead of the printer default (which is 0 OD for a reflective media printer).

Evaluation

The hardcopy should match the GSDF for Illumination=150, Reflected Ambient Light=0. The minimum density should be 0.4 OD. Use Excel spreadsheet described in section 2.4 to assess conformance. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #125: Collation (optional)

This test prints the dsitest pattern using an N-ACTION request on Basic Film Session level (instead of the Basic Film Box level which is used in all other tests).

Evaluation

The result should equal the result from test #101. If no hardcopy is printed, the Print Server does not support collation and the test is not passed. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #126: 8-bit Image Transmission (optional)

This test prints the dsitest pattern. The pattern is transmitted with 8 bits/pixel in the Basic Grayscale Image Box.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #127: 12-bit Image Transmission (optional)

This test prints the dsitest pattern. The pattern is transmitted with 12 bits/pixel in the Basic Grayscale Image Box.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #128: Pixel Aspect Ratio

This test prints the Pixel Aspect Ratio Test Pattern with a pixel aspect ratio of 4:1.

Evaluation

The hardcopy should display a square image. If the image is a rectangle, the Print Server does not support Pixel Aspect Ratio and the test is not passed.

Test #129: Requested Image Size (optional)

This test prints the dsitest pattern with a requested image size of 127mm (5”).

Evaluation

The hardcopy should display the dsitest image with a size of 127x127mm (5x5”). If no hardcopy is created, or if the image is printed with a size different from 127x127mm, the Print Server does not support Requested Image Size and the test is not passed. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #130: Polarity (optional)

This test sends an inverted dsitest pattern to the printer and sets polarity to “REVERSE”.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #131: MONOCHROME 1 (optional)

This test sends a dsitest pattern using MONOCHROME1 photometric interpretation to the printer.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

### Testing the Presentation LUT Implementation

This set of tests checks details of the Print SCPs implementation of the Presentation LUT SOP Class.

Test #140: LIN OD Presentation LUT Shape (optional)

This test sends a dsitest pattern that has been converted to a value space linear to optical density. A Presentation LUT Shape of “LIN OD” is used. The image has been processed to be linear to the optical densities of a hardcopy printed with Dmax = 3.0 OD, Dmin = 0.2 OD, on a GSDF calibrated printer with Illumination = 150 cd/m2 and Reflected Ambient Light = 0 cd/m2.

Evaluation

The result should equal the result from test #101, or at least be very similar. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #141: Presentation LUT with 256 entries, Image with 8 bits/pixel (optional)

This test print creates a Presentation LUT with 256 entries and transmits the dsitest pattern with 8 bits/pixel.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #142: Presentation LUT with 4096 entries, Image with 12 bits/pixel (optional)

This test print creates a Presentation LUT with 4096 entries and transmits the dsitest pattern with 12 bits/pixel.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #143: Presentation LUT with 256 entries, Image with 12 bits/pixel (optional)

This test print creates a Presentation LUT with 256 entries and transmits the dsitest pattern with 12 bits/pixel. Combinations of Presentation LUT and pixel data where the number of LUT entries does not match with the possible number of grayscale values in the pixel data have been declared a protocol violation in CP 202 (2001) but were allowed in earlier versions of the DICOM standard. Therefore, Print Servers may reject this test.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #144: Presentation LUT with 4096 entries, Image with 8 bits/pixel (optional)

This test print creates a Presentation LUT with 4096 entries and transmits the dsitest pattern with 8 bits/pixel. Combinations of Presentation LUT and pixel data where the number of LUT entries does not match with the possible number of grayscale values in the pixel data have been declared a protocol violation in CP 202 (2001) but were allowed in earlier versions of the DICOM standard. Therefore, Print Servers may reject this test.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #145: Presentation LUT with 256 entries and 11 bits/entry (optional)

This test print creates a Presentation LUT with 256 entries and transmits the dsitest pattern with 8 bits/pixel. The Presentation LUT is unusual because the number of significant bits per entry is 11.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #146: Presentation LUT with 256 entries and 13 bits/entry (optional)

This test print creates a Presentation LUT with 256 entries and transmits the dsitest pattern with 8 bits/pixel. The Presentation LUT is unusual because the number of significant bits per entry is 13.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #147: Scrambled Presentation LUT and Image (optional)

This test print creates a Presentation LUT with 4096 entries and transmits the dsitest pattern with 12 bits/pixel. Both Presentation LUT and Image pixel data are scrambled such that only a correct application of the Presentation LUT restores a visible image.

Evaluation

The result should equal the result from test #101. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #148: Presentation LUT with OW value representation (optional)

This test print creates a Presentation LUT with 4096 entries and transmits the dsitest pattern with 12 bits/pixel. The Presentation LUT is unusual because the DICOM Value Representation of the LUT is “OW”. This value representation has been introduced for LUT Data in DICOM Supplement 33 in 1999.

Evaluation

The log file for this test case (log/048.log) should be examined. If the Print Server accepts the Presentation LUT N-CREATE request, it may be assumed that the Print Server supports LUT Data with OW value representation. Otherwise, the test is not passed. This test is optional. Passing the test is not a requirement for IHE Print Servers.

### Testing Printer Layouts

This set of tests tries out the most common image display formats (page layouts). All tests in this section are optional.

Test #150: 2x2 Layout (optional)

This test creates and prints a Basic Film Box with an image display format (page layout) of “STANDARD\2,2”.

Evaluation

The result should show the dsitest and the wedge pattern tiled on a print-out with 2x2 image boxes. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #151: 3x5 Layout (optional)

This test creates and prints a Basic Film Box with an image display format (page layout) of “STANDARD\2,2”.

Evaluation

The result should show the dsitest and the wedge pattern tiled on a print-out with 3x5 image boxes. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #152: 4x5 Layout (optional)

This test creates and prints a Basic Film Box with an image display format (page layout) of “STANDARD\2,2”.

Evaluation

The result should show the dsitest and the wedge pattern tiled on a print-out with 4x5 image boxes. This test is optional. Passing the test is not a requirement for IHE Print Servers.

Test #153: 4x6 Layout (optional)

This test creates and prints a Basic Film Box with an image display format (page layout) of “STANDARD\2,2”.

Evaluation

The result should show the dsitest and the wedge pattern tiled on a print-out with 4x6 image boxes. This test is optional. Passing the test is not a requirement for IHE Print Servers.

## Using the Excel Spreadsheet to Assess GSDF Conformance

The Excel file “gsdf_od.xls” contains two tables: the first table is used to enter the measured optical density (OD) values and some administrative data, the second one is automatically filled and used for printing the results.  Figure 1 shows the “Input” table – the sheet is protected and only the blue fields can be modified.

Figure 1: “Input” table of the Excel file for hardcopy calibration

Note:  the spreadsheet contains macros which are required for operation, i. e. they should be activated if Excel asks the user whether to activate or deactivate macros while loading the file.

### Open Excel File and Fill Administrative Data

First you should open the Excel file “gsdf_od.xls” (activate macros if a requesting dialog appears) and enter the administrative data: Vendor, Device, Date, Time and the Performing Technician.

### Measure Boxes and Enter Results

The test case description for the test case being evaluated describes the values for Illumination and Reflected Ambient Light which are used for the test case.  These values must be filled in the Excel form.  The test case has produced a hardcopy showing the “dsitest” pattern (Figure 2).

Figure 2: Philips “dsitest” image

A densitometer should be used according to the description shipped with the device to measure the optical density for the 16 reference fields on the “dsitest pattern”.  The measured value (in OD) should be entered into the corresponding field in the Excel “Input” table (e. g. “measured Box #1” for the white box at the bottom left-hand corner).  The same procedure has to be performed for the all 16 boxes (see Figure 3 for the correct order).

Figure 3: Philips “dsitest” image (detail), order of boxes

After all 16 measured values have been entered the “overall Status” should be “OK”.  Otherwise you have most probably made a mistake, e. g. measured the fields in wrong order.

### Print “Output” Page and Assess Deviation from GSDF

After successfully measuring and entering the data the „Output“ page shows the result (see Figure 4).

Figure 4: Printed „Output“ table of the Excel file for hardcopy calibration

This page can be printed on standard paper.  Since all necessary information is contained on this page there is no need to print the “Input” table.  The table below the administrative data shows the measured optical density values versus the GSDF (converted to OD) and the deviation from the standard curve (absolute and percentage values as well as the mean difference and the standard deviation).  Unfortunately, no precise formula for an assessment of the deviation can be given.  Obviously, deviation from the GSDF should be as small as possible and in any case below 20%.  In addition, the curve should be “smooth” and should not show any steps.  This is the reason why both curves are shown in a diagram.  The 16 measured values are interpolated by a built-in algorithm – this has no influence on the original data presented in the table, though.

IHE Print Server Actors

Test Results

# Print Server Test Results

Print Server vendor

Name and software version of the Print Server

Person filling in this form, place, date

Print medium: transmissive (film) or reflective (paper)

## Testing the GSDF Calibration

| No. | Test Name | Test Result |
| --- | --- | --- |
| x01 | GSDF Curve under Default Lighting Conditions |  |
| x02 | GSDF Curve under unusual Lighting Conditions |  |
| x03 | GSDF Curve with Presentation LUT (256 Entries), Default Lighting Conditions |  |
| x04 | GSDF Curve with Presentation LUT (4096 Entries), Default Lighting Conditions |  |
| x05 | GSDF Curve with Presentation LUT (256 Entries), unusual Lighting Conditions |  |
| x06 | GSDF Curve with Presentation LUT (4096 Entries), unusual Lighting Conditions |  |

## Testing the Visual Image Quality

| No. | Test Name | Test Result |
| --- | --- | --- |
| x10 | Test for Visual Artifacts, 8-bit image data with Presentation LUT Shape |  |
| x11 | Test for Visual Artifacts, 8-bit image data with Presentation LUT (256 entries) |  |
| x13 | Test for Visual Artifacts, 12-bit image data with Presentation LUT Shape |  |
| x15 | Test for Visual Artifacts, 12-bit image data with Presentation LUT (4096 entries) |  |

## Testing the General Print Protocol Implementation

| No. | Test Name | Test Result |
| --- | --- | --- |
| x20 | Referenced Presentation LUT SQ on Film Session Level (optional) |  |
| x21 | Referenced Presentation LUT SQ on Film Box Level |  |
| x22 | Landscape Orientation |  |
| x23 | Maximum Density |  |
| x24 | Minimum Density (optional) |  |
| x25 | Collation (optional) |  |
| x26 | 8-bit Image Transmission (optional) |  |
| x27 | 12-bit Image Transmission (optional) |  |
| x28 | Pixel Aspect Ratio |  |
| x29 | Requested Image Size (optional) |  |
| x30 | Polarity (optional) |  |
| x31 | MONOCHROME 1 (optional) |  |

## Testing the Presentation LUT Implementation

| No. | Test Name | Test Result |
| --- | --- | --- |
| x40 | LIN OD Presentation LUT Shape (optional) |  |
| x41 | Presentation LUT with 256 entries, Image with 8 bits/pixel (optional) |  |
| x42 | Presentation LUT with 4096 entries, Image with 12 bits/pixel (optional) |  |
| x43 | Presentation LUT with 256 entries, Image with 12 bits/pixel (optional) |  |
| x44 | Presentation LUT with 4096 entries, Image with 8 bits/pixel (optional) |  |
| x45 | Presentation LUT with 256 entries and 11 bits/entry (optional) |  |
| x46 | Presentation LUT with 256 entries and 13 bits/entry (optional) |  |
| x47 | Scrambled Presentation LUT and Image (optional) |  |
| x48 | Presentation LUT with OW value representation (optional) |  |

## Testing the Presentation LUT Implementation

| No. | Test Name | Test Result |
| --- | --- | --- |
| x50 | 2x2 Layout (optional) |  |
| x51 | 3x5 Layout (optional) |  |
| x52 | 4x5 Layout (optional) |  |
| x53 | 4x6 Layout (optional) |  |

## General Comments

Please fill in general comments to the “Print Server Test”.

General Comments

# References

[1]	NEMA Standards Publications PS 3.14-1999, Digital Imaging and Communications in Medicine (DICOM), Part 14: Grayscale Standard Display Function, National Electrical Manufacturers Association, Rosslyn, VA, 1999.

[2]	DICOM Standards Committee, Working Group 6 Printer Ad Hoc: Digital Imaging and Communications in Medicine (DICOM), Supplement 22: Presentation Look Up Table (LUT), Final Text, January 1998.