Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Association Services  -  Association Abort

Level_2 Test-Case Specification Template

Document ID:		1.2.840.113702.1.2.7.1

Version Date:		10/20/95

Application Entities Interconnectivity

Association Services  -  Association Abort

Level_2 Test-Case Specification Template

i)	Purpose:

To define a test case identified by a test design specification

ii)	Outline:

A test-case specification shall have the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1.	Test-Case-Specification Identifier

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| A-ABORT | Abort an Association | Specified by testing team |

2.	Test Items

The test items involved are the two Application Entities being tested for interconnectivity.

For each test item, reference the following documentation:

Requirements specification

Design specification

Users guide

Operations guide

Installation guide

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statement of each Application Entity to be tested.

3.	Input Specifications

Specify each input required to execute the test case, by value, by filename, etc.

Identify all appropriate data bases, files, terminal messages, memory resident areas, etc.

Specify all relationships between inputs.

The input specifications required for testing the A-ABORT Upper Layer service shall be the user interaction required to cause the requestor Application Entity to issue an A-ABORT request.

The assumptions for conducting this test are as follows:

1.	the ability to invoke the A-ABORT service is accessible to the operator

2.	the A-ASSOCIATE service has previously been successfully invoked.

4.	Output Specifications

Specify all the outputs and features required of the test items.

The output and features expected of the Application Entities for successful execution of the A-ABORT service are as follows:

Output:

Log files entries:

Event Level Information:

| Requestor | Service Primitive Direction |  | Acceptor |
| --- | --- | --- | --- |
| A-ABORT request |  |  |  |
|  |  |  | A-ABORT indication |

Features:

At this point, the association between the two Application Entities has been released and no further communication is allowed.

5.	Environmental Needs

5.1		Hardware

The computer system required for each of the Application Entities is as defined by the respective suppliers of the Application Entities.

The assumption is that the systems are networked together and therefore the appropriate hardware connections to the respective networks are required.

5.2		Software

The system and application software required for each of the Application Entities is as defined by the respective suppliers of the Application Entities.

The defined communication modes for DICOM are: 1) OSI,  2) TCP/IP,  3) Point-to-Point. The assumption is that each system has the appropriate communication software available to the Application Entiry.

5.3		Other environmental requirements

6.	Special Procedural Requirements

7.	Intercase Dependencies

The A-ASSOCIATE service must be tested successfully before the A-ABORT service can be tested.

In performing the test for the A-ABORT service, the A-ASSOCIATE must be successfully invoked before the A-ABORT service can be invoked.