Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Abort Transfer

Document ID:		1.2.840.113702.1.3.18.1

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Abort Transfer

i)	Purpose:

This Test-Case Specification document defines a test of the behavior of the General Electric Advantage Review workstation (CT Storage SCP) when an association is aborted by the SCU.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1. Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Abort Transfer | Remote abort of a transfer | 1.2.840.113702.1.3.18.1 |

2. Test Items

The behavior of the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation will be tested under the condition of a transfer abort issued by the SCU.

3. Input Specifications

An in-process transfer is aborted by the SCU when it is approximately 1/2 complete.

4. Output Specifications

Monitor and record the status display which must indicate transfer aborted.

5. Environmental Needs

5.1	Hardware

A GE Advantage Review Workstation networked together with a system that acts as a CT storage SCU.

5.2	Software

Advantage Review Software Revision 1.03.

5.3	Other environmental requirements

A telephone must be accessible to the operator of the workstation and of the SCU so that transfer start events can be communicated

6. Special Procedural Requirements

The status display on the Advantage Review workstation updates rapidly and overwrites the previous message requiring close scrutiny of the display.

7. Intercase Dependencies

If the association is not cleanly and completely aborted, there may be an impact on subsequent association establishment.