Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Parallel Associations Boundary

Document ID:		1.2.840.113702.1.3.10.2

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Parallel Associations Boundary

i)	Purpose:

This Test-Case Specification document defines a test of boundary condition behavior in Multiple Association Mode, for the Picker PQ2000 CT scanner (CT Storage SCU). In Multiple Association Mode the storage SCU is capable of transmitting streams of images over multiple simultaneous associations. Studies for different patients will be transmitted simultaneously to achieve a large number of parallel associations.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1.	Test-Case-Specification Identifier

This example Test Case Specification has the ID

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Parallel Associations Boundary | CT Image Storage Services in Multiple Association Mode for Multiple Studies to Maximize Parallel Associations | 1.2.840.113702.1.3.10.2 |

2.	Test Items

The ability of the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner to support a large number of parallel associations will be tested in Multiple Association Mode.

3.	Input Specifications

Using the netconfig tool set the configuration to Multiple Association Mode.

It is not possible to control the number of simultaneous associations established by the HANI Application Entity as this depends on internal queuing and SCP negotiation.

Select 8 CT exams with approximately 20 images each.

At the PQ2000 operator's console select HyperLan transfer destination, select one case each for n Patients from the directory and initiate transfer.

4.	Output Specifications

Monitor and record the status display which must indicate transfer complete  (HyperLan Network Transfer Complete).

5.	Environmental Needs

5.1		Hardware

A Picker PQ2000 CT scanner operator's console networked together with a system that acts as a CT storage SCP.

5.2		Software

Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner.

5.3		Other environmental requirements

A telephone must be accessible to the operator of the CT Scanner console and of the SCP so that transfer start events can be communicated

6.	Special Procedural Requirements

Two trained test personnel are required.

7.	Intercase Dependencies

None.