Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Parallel Associations Boundary

Document ID:		1.2.840.113702.1.3.19.1

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Parallel Associations Boundary

i)	Purpose:

This Test-Case Specification document defines a test of boundary condition behavior  of the General Electric Advantage Review workstation (CT Storage SCP). Studies for different patients will be transmitted simultaneously in an attempt to achieve a large number of parallel associations.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1. Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Parallel Associations Boundary | CT Image Storage Services for Multiple Studies over Parallel Associations | 1.2.840.113702.1.3.19.1 |

2. Test Items

The ability of the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation to support a large number of parallel associations will be tested.

3. Input Specifications

From the SCU transmit 8 CT exams simultaneously over multiple parallel associations. The best SCU implementation would create at least 8 parallel associations.

4. Output Specifications

Monitor and record the status display and observe the patient directory upon completion of transfer. The status display must indicate transfer complete and the study should appear in the directory.

5. Environmental Needs

5.1	Hardware

A GE Advantage Review Workstation networked together with a system that acts as a CT storage SCU.

5.2	Software

Advantage Review Software Revision 1.03.

5.3	Other environmental requirements

A telephone must be accessible to the operator of the Advantage Review workstation and of the SCU so that transfer start events can be communicated.

6. Special Procedural Requirements

The status display on the Advantage Review workstation updates rapidly and overwrites the previous message requiring close scrutiny of the display.

7. Intercase Dependencies

The available disk storage on the Advantage Review workstation must be monitored after each transfer to avoid resource limit errors.