Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

SCP Disk Full Resource Boundary

Document ID:		1.2.840.113702.1.3.20.1

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

SCP Disk Full Resource Boundary

i)	Purpose:

This Test-Case Specification document defines a test of boundary condition behavior  of the General Electric Advantage Review workstation (CT Storage SCP). Association establishment and CT image transfer will be attempted when the local storage system is full.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1. Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| SCP Disk Full Boundary | CT Image Storage Services attempted transfer with SCP Disk Full | 1.2.840.113702.1.3.20.1 |

2. Test Items

The ability of the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation to correctly detect and report an SCP Disk Full error state will be tested.

3. Input Specifications

From the SCU transmit a CT exam with approximately 20 images.

4. Output Specifications

Monitor and record the status display. A resource error must be reported and the transfer aborted with a proper status returned to the SCU.

5. Environmental Needs

5.1	Hardware

A GE Advantage Review Workstation networked together with a system that acts as a CT storage SCU.

5.2	Software

Advantage Review Software Revision 1.03.

5.3	Other environmental requirements

A telephone must be accessible to the operator of the Advantage Review workstation and of the SCU so that transfer start events can be communicated.

6. Special Procedural Requirements

The status display on the Advantage Review workstation updates rapidly and overwrites the previous message requiring close scrutiny of the display.

7. Intercase Dependencies

The available disk storage on the Advantage Review workstation will be exhausted by this test.