Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Repeat Transfer

Document ID:		1.2.840.113702.1.3.17.1

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Repeat Transfer

i)	Purpose:

This Test-Case Specification document defines a test of the behavior of the General Electric Advantage Review workstation (CT Storage SCP) when a successful transfer is repeated.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1. Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Repeat Transfer | Receive a study twice | 1.2.840.113702.1.3.17.1 |

2. Test Items

The behavior of the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation will be tested under the condition of repeated transfer of a set of images.

3. Input Specifications

The SCU will transmit a set of images (study) for a selected patient (about 20 images). Upon successfully completing this transfer, the same set of images will be transferred again.

4. Output Specifications

Monitor and record the status display and observe the patient directory upon completion of transfer. For the first transfer the status display must indicate transfer complete and the study should appear in the directory. For the second (repeated) transfer, the status display should indicate the images and study are already present and that a warning has been returned for each image.

5. Environmental Needs

5.1	Hardware

A GE Advantage Review Workstation networked together with a system that acts as a CT storage SCU.

5.2	Software

Advantage Review Software Revision 1.03.

5.3	Other environmental requirements

A telephone must be accessible to the operator of the Advantage Review workstation and of the SCU so that transfer start events can be communicated.

6. Special Procedural Requirements

The status display on the Advantage Review workstation updates rapidly and overwrites the previous message requiring close scrutiny of the display.

7. Intercase Dependencies

The available disk storage on the Advantage Review workstation must be monitored after each transfer to avoid resource limit errors.