Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Association Establishment and Image Transfer

Single Association Mode, Multiple Study

Level_3 Test-Case Specification Example

Document ID:		1.2.840.113702.1.3.5.1

Version Date:		11/12/95

Application Entities Interconnectivity

Association Establishment and Image Transfer

Single Association Mode, Multiple Study

Level_3 Test-Case Specification Example

i)	Purpose:

This Test-Case Specification document defines a test of basic DICOM Association Establishment and CT image transfer in Single Association Mode, between the Picker PQ2000 CT scanner (CT Storage SCU) and the General Electric Advantage Review workstation (CT Storage SCP). Multiple studies for one patient and studies for several different patients will be exchanged.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1.	Test-Case-Specification Identifier

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Association Establishment and Image Transfer | Association establishment and CT Image Storage Services in Single Association Mode for Multiple Studies | 1.2.840.113702.1.3.5.1 |

2.	Test Items

Basic Association Establishment and CT image storage will be tested between the two Application Entities: the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner and the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation.  Multiple studies will be transferred.

3.	Input Specifications

Picker PQ2000 inputs:  Current CT exams for the following patients

| Patient Name | Exam | Number of Images |
| --- | --- | --- |
| Patient 1 | 1 |  |
| Patient 2 | 1 |  |
| Patient 3 | 1 |  |
| Patient 4 | 1 |  |
|  | 2 |  |

The user interaction input step at the Service Class User Application Entity is that required to:

1.	cause the Application Entity to issue a DIMSE C-STORE  request for the Service Object Pair under test.

At the PQ2000 operatorÕs console, select HyperLan transfer destination, select Patient 1 from the directory and initiate transfer.

4.	Output Specifications

At the SCU,  monitor and record the status display.

At the SCP, monitor and record the status display and observe the patient directory upon completion of transfer

5.	Environmental Needs

5.1		Hardware

The computer system required for each of the Application Entities is as defined by the respective suppliers of the Application Entities.

The assumption is that the systems are networked together and therefore the appropriate hardware connections to the respective networks are required.

5.2		Software

The system and application software required for each of the Application Entities is as defined by the respective suppliers of the Application Entities.

5.3		Other environmental requirements

6.	Special Procedural Requirements

Two trained test personnel are required.  The GE workstation status display updates rapidly overwriting previous messages. It must be monitored very carefully.

7.	Intercase Dependencies

The available disk storage on the SCP system must be monitored after each transfer to avoid resource limit errors.