Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Association Establishment and Image Transfer

Single Association Mode, Multiple Study

Document ID:		1.2.840.113702.1.3.5.2

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Association Establishment and Image Transfer

Single Association Mode, Multiple Study

i)	Purpose:

This Test-Case Specification document defines a test of basic DICOM Association Establishment and CT image transfer in Single Association Mode the Picker PQ2000 CT scanner (CT Storage SCU). Multiple studies for one patient and studies for several different patients will be transmitted

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1. Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Association Establishment and Image Transfer | Association establishment and CT Image Storage Services in Single Association Mode for Multiple Studies | 1.2.840.113702.1.3.5.2 |

2. Test Items

Basic Association Establishment and CT image storage functions of HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner will be tested.  Multiple studies will be transferred.

3. Input Specifications

Select exams of the type specified in the following table and transmit on separate associations.

| Patient Name | Exam | Number of Images |
| --- | --- | --- |
| Patient 1 | 1 | 40 |
| Patient 2 | 1 | 40 |
| Patient 3 | 1 | 40 |
| Patient 4 | 1 | 20 |
|  | 2 | 20 |

The user interaction input step at the PQ2000 console is that required to cause the Application Entity to issue a DIMSE C-STORE  request for the Service Object Pair under test.

At the PQ2000 operatorÕs console select HyperLan transfer destination, select study m for Patient n from the directory and initiate transfer.

4. Output Specifications

Monitor and record the status display which must indicate transfer complete (HyperLan Network Transfer Complete).

5. Environmental Needs

5.1	Hardware

A Picker PQ2000 CT scanner operatorÕs console networked together with a system that acts as a CT storage SCP.

5.2	Software

Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner

5.3	Other environmental requirements

A telephone must be accessible to the operator of the CT Scanner console and of the SCP so that transfer start events can be communicated

6. Special Procedural Requirements

Two trained test personnel are required.

7. Intercase Dependencies

None.