Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Association Establishment and Image Transfer, Multiple Study

Document ID:		1.2.840.113702.1.3.16.1

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Association Establishment and Image Transfer, Multiple Study

i)	Purpose:

This Test-Case Specification document defines a test of basic DICOM Association Establishment and CT image transfer for the General Electric Advantage Review workstation (CT Storage SCP). Multiple studies for one patient and studies for several different patients will be exchanged.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1. Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Association Establishment and Image Transfer | Association establishment and CT Image Storage Services for Multiple Studies | 1.2.840.113702.1.3.16.1 |

2. Test Items

Basic Association Establishment and CT image storage functions of the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation will be tested.  Multiple studies will be transferred.

3. Input Specifications

From the SCU transmit the following CT exams:

| Patient Name | Exam | Approximate Number of Images |
| --- | --- | --- |
| Patient 1 | 1 | 40 |
| Patient 2 | 1 | 40 |
| Patient 3 | 1 | 40 |
| Patient 4 | 1 | 20 |
| Patient 4 | 2 | 20 |

Each exam should be transmitted using a separate association.  Associations should be established and released sequentially.

4. Output Specifications

Monitor and record the status display and observe the patient directory upon completion of transfer. The status display must indicate transfer complete and the study should appear in the directory.

5. Environmental Needs

5.1	Hardware

A GE Advantage Review Workstation networked together with a system that acts as a CT storage SCU.

5.2	Software

Advantage Review Software Revision 1.03.

5.3	Other environmental requirements

A telephone must be accessible to the operator of the Advantage Review workstation and of the SCU so that transfer start events can be communicated.

6. Special Procedural Requirements

The status display on the Advantage Review workstation updates rapidly and overwrites the previous message requiring close scrutiny of the display.

7. Intercase Dependencies

The available disk storage on the Advantage Review workstation must be monitored after each transfer to avoid resource limit errors.