Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Association Establishment and Image Transfer

Single Association Mode, Single Study

Document ID:		1.2.840.113702.1.3.4.2

Version Date:		1/29/96

Application Entities Interconnectivity

Level_3 Test-Case Specification Example

Association Establishment and Image Transfer

Single Association Mode, Single Study

i)	Purpose:

This Test-Case Specification document defines a test of basic DICOM Association Establishment and CT image transfer in Single Association Mode, for the Picker PQ2000 CT scanner (CT Storage SCU).

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1.	Test-Case-Specification Identifier

This example Test Case Specification has the ID

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Association Establishment and Image Transfer | Association establishment and CT Image Storage Services in Single Association Mode for a Single Study | 1.2.840.113702.1.3.4.2 |

2.	Test Items

Basic Association Establishment and CT image storage functions of the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner will be tested.

3.	Input Specifications

Select a current CT exam (reference name - Patient 1) with approximately 20 images. The user interaction input step at the PQ2000 console is that required to cause the Application Entity to issue a DIMSE C-STORE  request for the Service Object Pair under test.

At the PQ2000 operatorÕs console, select HyperLan transfer destination, select Patient 1 from the directory and initiate transfer.

4.	Output Specifications

Monitor and record the status display which must indicate transfer complete  (HyperLan Network Transfer Complete).

5.	Environmental Needs

5.1		Hardware

A Picker PQ2000 CT scanner operator's console networked together with a system that acts as a CT storage SCP.

5.2		Software

Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner.

5.3		Other environmental requirements

A telephone must be accessible to the operator of the CT Scanner console and of the SCP so that transfer start events can be communicated.

6.	Special Procedural Requirements

Two trained test personnel are required.

7.	Intercase Dependencies

None.