Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Storage Service Class

DIMSE C-STORE Service Object Pair

Level_2 Test-Case Specification Template

Document ID:		1.2.840.113702.1.2.10.1

Version Date:		10/20/95

Application Entities Interconnectivity

Storage Service Class

DIMSE C-STORE Service Object Pair

Level_2 Test-Case Specification Template

i)	Purpose:

The purpose of a Test-Case Specification document is to define a test case identified by a test design specification

ii)	Outline:

A test-case specification shall have the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1.	Test-Case-Specification Identifier

The test-case-specification identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| C-STORE; Service Object Pair | Storage Service for Service Object Pair | Specified by testing team |

2.	Test Items

The test items involved are the two Application Entities being tested for inter-operability.

For each test item, reference the following documentation:

Requirements specification

Design specification

Users guide

Operations guide

Installation guide

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statement of each Application Entity to be tested.

3.	Input Specifications

Specify each input required to execute the test case, by value, by filename, etc.

Identify all appropriate data bases, files, terminal messages, memory resident areas, etc.

Specify all relationships between inputs.

The input specifications required for testing the DIMSE C-STORE service shall be the user interaction required to cause the Service Class User Application Entity to issue a C-STORE request for the Service Object Pair under test.

Documented as part of the input specification for the test of the Service Object Pair Class shall be the following items:

Application Context

Presentation Contexts

User Information Items

Service Object Pair Class

The user interaction input step at the Service Class User Application Entity is that required to:

1.	cause the Application Entity to issue a DIMSE C-STORE  request for the Service Object Pair under test.

4.	Output Specifications

Specify all the outputs and features required of the test items.

The output and features expected of the Application Entities for successful execution of the C-STORE service fall into three lavels or categories:

1.	Command Event Level

2.	Command Data Level

3.	Data Set Level

Output:

Log files entries:

Event Level Information:

| Service Class User | Service Primitive Direction |  | Service Class Provider |
| --- | --- | --- | --- |
| C-STORE request |  |  |  |
|  |  |  | C-STORE indication |
|  |  |  | C-STORE response |
| C-STORE confirmation |  |  |  |

Command Data Level Information:

The critical output information at the Command Data Level is the status information sent back to the Service Class User Application Entity in which the Service Class Provider Application Entity  indicates the result of the service. The C-STORE status choices, defined by the DICOM Standard PS3, are given in the following table.

| C-STORE Status |  |
| --- | --- |
| Service Status | Further Meaning |
| Refused | Out of Resources |
|  | Data Set does not match SOP Class |
| Error | Cannot Understand |
|  | Coercion of Data Elements |
| Warning | Data Set does not match SOP Class |
|  | Elements discarded |
| Success |  |

Data Set Level Information:

Features:

5.	Environmental Needs

5.1		Hardware

The computer system required for each of the Application Entities is as defined by the respective suppliers of the Application Entities.

The assumption is that the systems are networked together and therefore the appropriate hardware connections to the respective networks are required.

5.2		Software

The system and application software required for each of the Application Entities is as defined by the respective suppliers of the Application Entities.

The defined communication modes for DICOM are: 1) OSI,  2) TCP/IP,  3) Point-to-Point. The assumption is that each system has the appropriate communication software available to the Application Entiry.

5.3		Other environmental requirements

6.	Special Procedural Requirements

7.	Intercase Dependencies