Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Basic Network Connectivity

Level_3 Test-Case Specification Example

Document ID:		1.2.840.113702.1.3.14.1

Version Date:		1/29/96

Application Entities Interconnectivity

Basic Network Connectivity

Level_3 Test-Case Specification Example

i)	Purpose:

This Test-Case Specification document defines a test of basic TCP/IP connectivity for the General Electric Advantage Review workstation as installed at the Milton S. Hershey Medical Center, Hershey, PA.

ii)	Outline:

This test-case specification has the following structure:

1.	Test-case-specification identifier

2.	Test items

3.	Input specification

4.	Output specification

5.	Environmental needs

6.	Special procedural requirements

7.	Intercase dependencies

1.	Test-Case-Specification Identifier

This example Test Case Specification has the ID:

| Test Case | Test Case Description | Test-Case-Specification Identifier |
| --- | --- | --- |
| Basic Network Connectivity | TCP/IP communication is possible between the system under test and an independent test node | 1.2.840.113702.1.3.14.1 |

2.	Test Items

Basic network connectivity of the General Electric Advantage Review workstation running Software Revision 1.03 will be tested. The extent to which connectivity can be assessed is dependent on the available of TCP/IP based applications in the system

3.	Input Specifications

GE Advantage Review inputs: ping pqeast.xray.hmc.psu.edu.

4.	Output Specifications

Correct response from ping: system is alive.

5.	Environmental Needs

5.1		Hardware

The assumption is that the systems are networked together and therefore the appropriate hardware connections to the respective networks are required.

5.2		Software

The ping utility is required on both systems. Other TCP based functions such as telnet and ftp could also be use, but were not available on the PQ 2000.

5.3		Other environmental requirements

6.	Special Procedural Requirements

None.

7.	Intercase Dependencies

TCP/IP connectivity is a fundamental requirement for all other DICOM interconnectivity tests. The systems under test were attached to the campus backbone of the Milton S. Hershey Medical Center and Penn State College of Medicine and so were dependent on institutional resources for routing, name services and other TCP/IP functions.