Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Association Services

Level_2 Test-Design Specification Template

Document ID:		1.2.840.113702.1.2.3.1

Version Date:		10/20/95

Application Entities Interconnectitivy

Association Services

Level_2 Test-Design Specification Template

i)	Purpose

The purpose of the Test Design Specification document is to detail refinements of the test approach presented in the Test Plan.

To identify the features to be tested by the design and its associated tests.

ii)	Outline

A test-design specification document shall have the following structure.

1.	Test-design-specification identifier

2.	Features to be tested

3.	Approach refinements

4.	Test identification

5.	Feature pass/fail criteria

1.	Test-Design-Specification Identifier

The Test-Design-Specification Identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

2.	Features to be Tested

Identify test items and describe the features and combination of features which are the object of this design specification.

For each feature or features combination, a reference to its associated requirements in the item requirement specification or design description should be included.

The following features will be tested in every test of interconnectivity of the Association Services between two Application Entities.

Association Services:

A-ASSOCIATE		Establish an Association

A-RELEASE		Release an Association

A-ABORT		Abort an Association

C-ECHO		Test for connectivity

P-DATA		Send a DICOM message

Association Service Provider Internal Integrity:

A-P-ABORT

Testing guidelines for the Association Services will be described in detail at this level of the Testing Guidelines documents. Testing guidelines for the Service Classes and for the internal integrity of the Association Service Provider will be described in appropriate documents at lower levels in the tree of guideline documents.

The documentation associated with each test should include:

Test Design Specification Number

Identification of the Conformanace Statements with specific reference to the relevant sections in the Conformance Statements most concerned with the specific test.

Identification of other relevant documentation of the Application Entities with specific reference to the relevant sections of the documentation most concerned with the specific test.

3.	Approach Refinements

Specify refinements to the approach described in the test plan.

Include specific test techniques to be used.

Identify method of analyzing test results.

The approach used for testing the Association Services involves a complete exercising of the services. The assumptions are:

1.	The services will be invoked from the user interface level of operation

2.	Each invocation of the Association Service primitives is logged

3.	Information about the completion of an Association Service, successful or unsucessful, may or may not be passed to the user interface level of operation

The actual design of tests will depend on the the amount of information and the control that the designers and implementers have made available to the system operators. In testing for interconnectivity at the Association Service level, there are two levels of information access that may be available to the user:

Event Access:	Access to the information that a significant event has occurred

Data Access:	Access to the information that a significant event has occurred and access to the data accompanying the event.

The assumption to be made in this template version of the test design is that at least Event Access level of information is available for testing. The approach taken for each test will be to run the test and to ensure that the significant events have taken place in the proper sequence. If the Data Access level of information is available, then the data can be checked more closely.

The primary method of analyzing test results will be by examining ‘log’ files, both at the application level of association negotiation and at the Upper Layer Association Services level.

4.	Test Identification

List the identifier and a brief description of each test case associated with this design.

Each Test Identifier shall be a unique identifier for both providers of the Application Entities to be tested.

At this level of testing, the following set of Association Services will be tested.

Association Services:

A-ASSOCIATE		Establish an Association

A-RELEASE		Release an Association

A-ABORT		Abort an Association

P-DATA		Send a DICOM message

5.	Feature Pass/Fail Criteria

The Pass/Fail criteria for testing the Association Services at this level are expressed in terms of the events that can be monitored and that can be expected to be logged in the log files of the Application Entities and the Upper Layer Association Services.