Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_4 Test Incident Report Example

Document ID:		1.2.840.113702.1.4.5.1

Version Date:		11/12/96

Applications Entities Interconnectivity

Level_4 Test Incident Report Example

i.	Purpose

The purpose of the Test Incident Report is to document anomalous events that occurred during interconnectivity testing between a Picker PQ2000 CT scanner and a General Electric Advantage Review Workstation.

ii.	Outline

This test-incident report has the following structure

(1)	Test-incident report identifier

(2)	Summary

(3)	Incident description

(4)	Impact

1.	Test-Incident Report Identifier

The Test Incident Report Identifier shall be a unique identifier for both suppliers of the Application Entities being tested. This Test Incident Report Example has the ID: 1.2.840.113702.1.4.5.1.

2.	Summary

The items tested for interconnectivity were the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation and  the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner. The HANI AE served as the SCU and the Local DICOM Import AE served as the SCP. A Joint Test Schema (ID: 1.2.840.113702.1.4.1.1) was executed on November 3-4, 1995 by Fred Prior, Ph.D. and Franklyn Bradshaw, Ph.D.  Test logs were generated for each test item:

Test Log ID: 1.2.840.113702.1.4.2.1  - HANI Application Entity

Test Log ID: 1.2.840.113702.1.4.3.1  - Local DICOM Import Application Entity.

Three of the nine test cases specified in the Joint Test Schema produced anomalous results.

3.	Incident Description

| Incident | Test Case | Inputs | Expected Results | Anomalous Results | Observations |
| --- | --- | --- | --- | --- | --- |
| 1 | Repeatability of Transfers | Data Set 6: 9 images | Warning or error returned to SCU | Both Transfers successful | Both copies were stored by the SCP |
| 2 | Association Abort | Data Set 7: 28 images | Abort error message and interruption of transfer | Transfer aborted after image 2 on SCU but SCP received 15 images | Transfers queued, abort only takes effect after queue is emptied |
| 3 | Boundary Behavior - Resource boundary | Data Set 21: 44 images | Disk Full error reported by SCP to SCU | Transfer completed successfully | Images were discarded after receipt |

4.	Impact