Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Level_4 Test Report Example, Test Log, Example System 2

Document ID:		1.2.840.113702.1.4.3.1

Version Date:		6/12/96

Applications Entities Interconnectivity

Level_4 Test Report Example

Test Log, Example System 2

i.	Purpose

To provide a chronological record of relevant details about the execution of interconnectivity tests between a Picker PQ2000 CT scanner and a General Electric Advantage Review Workstation.  This log represents the results of testing of the General Electric Advantage Review Workstation.

ii.	Outline

A test log shall have the following structure:

(1)	Test log identifier

(2)	Description

(3)	Activity and event entries

1.	Test-Log Identifier

The Test Log Identifier shall be a unique identifier for both suppliers of the Application Entities tested.  This Test Log Example has the ID: 1.2.840.113702.1.4.3.1.

2.	Description

The items to be tested for interconnectivity are the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the General Electric Advantage Review Workstation and  the HANI Application Entity as implemented in Software Revisions 4.2 and 4.3A of the Picker PQ 2000 CT Scanner.

A GE Advantage Review Workstation which is in routine clinical use was utilized for testing after normal working hours. Similarly, the Picker PQ2000 CT scanner was temporarily taken out of service and dedicated for testing. The Advantage Review workstation was connected via level 5 Unshielded Twisted Pair to a Lantronix 10BaseT wiring concentrator/repeater. The Picker CT scanner was connected via 10Base5 coaxial cable to a Cabletron Mini-MMAC configured as a bridge. Both the Lantronix and Cabletron units were connected via multi-mode optical fiber to two different ports on a Cisco 2500 series enterprise router. This router is part of the campus backbone of the Milton S. Hershey Medical Center and Penn State College of Medicine. The network connecting the two devices under test experienced representative (albeit off hours) medical center network traffic during the testing period.

3.	Activity and Event Entries

3.1	Execution Description

The Joint Test Schema (ID: 1.2.840.113702.1.4.1.1) was executed on November 3-4, 1995 by Fred Prior, Ph.D. and Franklyn Bradshaw, Ph.D.

3.2	Procedure Results

| Test Case | Test Case ID | Results | P/F |
| --- | --- | --- | --- |
| Basic TCP/IP connectivity | 1.2.840.113702.1.3.14.1 | pqeast is alive | Pass |
| Association Establishment and image transfer (SCU and SCP roles) | 1.2.840.113702.1.3.15.1 | Data Set 1: 26 images; ÒDICOM Import Received : 26 Images from HANIÓ | Pass |
| Association Establishment - Single Association Mode | 1.2.840.113702.1.3.16.1 | Data Sets 2-5: 55, 49, 45, 48 images; ÒDICOM Import Received : {55, 49, 45, 48} Images from HANIÓ | Pass |
| Repeatability of Transfers | 1.2.840.113702.1.3.17.1 | Data Set 6: 9 images; ÒDICOM Import Received : 9 Images from HANIÓ; study appeared twice in the local directory. | Pass ? |
| Association Abort | 1.2.840.113702.1.3.18.1 | Data Set 7: 28 images; ÒDICOM Import Received : 15 Images from HANIÓ | Pass ? |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.15.1 | Data Set 8: 44 images; ÒDICOM Import Received : 44 Images from HANIÓ | Pass |
| Association Establishment - Multiple Association Mode | 1.2.840.113702.1.3.16.1 | Data Sets 9-12: 61, 58, 70, 65 images; ÒDICOM Import Received : {61, 58, 70, 65} Images from HANIÓ | Pass |
| Boundary Behavior - Multiple Simultaneous Associations | 1.2.840.113702.1.3.19.1 | Data Sets 13-20: approx. 60 images each; ÒDICOM Import Received : {n} Images from HANIÓ | Pass |
| Boundary Behavior - Resource boundary | 1.2.840.113702.1.3.20.1 | Data Set 21: 44 images; ÒDICOM Import Received : 44 Images from HANIÓ; study did not appear in the local directory. | ? |

3.3	Environmental Information

3.4	Anomalous Events

It was anticipated that the Repeatability of Transfers Test Case (ID: 1.2.840.113702.1.3.6.2) would generate an error (warning) condition. It did not. The same case was transmitted twice and accepted by the SCP both times.

The Association Abort test yielded an anomalous result. The SCU aborted after 2 images were successfully transferred, however the SCP received 15 images. Twenty eight images were originally queued for transfer.

The Disk Full boundary test yielded an anomalous result. It was assumed that an error status (warning) would be returned indicating the disk full condition. A success status was returned but the images were not inserted into the local directory.

3.5	Incident-Report Identifiers

A joint Incident Report covering both SCU and SCP behavior was generated (ID: 1.2.840.113702.1.4.5.1).