Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Storage Service Class

Level_2 Test Plan Template

Document ID:		1.2.840.113702.1.2.2.1

Version Date:		10/20/95

Application Entities Interconnectivity

Storage Service Class

Level_2 Test Plan Template

i.	Purpose.

The purpose of a Test Plan document is to prescribe the scope, approach, resources, and schedule of the testing activities.

To identify the items being tested, the features to be tested, the testing tasks to be performed, the personnel responsible for each task, and the risks associated with the test plan.

This document is the Level_2 Test Plan template document for the Storage Service Class and is intended to be the template base document from which all Test Plan documents for the Storage Service Class are derived. This document is intended to represent a generic Test Plan which accomodates all of the characteristics of the Storage Service Class for inter-operating Application Entities that follow the DICOM Standard, ACR/NEMA Standards Publication No. PS 3.

ii.	Outline.

A test plan document shall have the following structure:

1.	Test-plan identifier

2.	Introduction

3.	Test items

4.	Features to be tested

5.	Features not to be tested

6.	Approach

7.	Item pass/fail criteria

8.	Suspension criteria and resumption requirements

9.	Test deliverables

10.	Testing tasks

11.	Environmental needs

12.	Responsibilities

13.	Staffing and training needs

14.	Schedule

15.	Risks and contingencies

16.	Approvals

1.	Test-Plan Identifier

The Test Plan Identifier shall be a unique identifier for both suppliers of the Application Entities to be tested.

2.	Introduction

2.1	Objectives.  A test plan for testing the interconnectivity between two Application Entities shall support the following objectives.

(1)  To detail the activities required to prepare for and support the test.

(2)  To communicate to all responsible parties the tasks which they are to perform and the schedule to be followed in perfoming the tasks.

(3)  To define the sources of information used to prepare the test plan.

(4)  To define the test tools and environment needed to conduct the test.

2.2	Background.  The DICOM Standard, ACR/NEMA Standards Publication No. PS 3 has been developed to define a standard method of transferring images and associated information between devices maunfactured by different vendors. The standard defines various information objects and the operations or services that may be performed on the objects, organizes the objects and services into Service-Object Pair (SOP) Classes, and groups the SOP Classes into Service Classes which can be supported by any Application Entity. The DICOM Standard also allows each implementor of an Application Entity to determine the degree to which the Application Entity will conform to the DICOM Standard for each Service-Object Pair Class.

2.3	Scope.  A test for interconnectivity between two Application Entities has the objective to show that the Application Entities are able to exchange messages in accordance with the DICOM Standard, ACR/NEMA Standards Publication No. PS 3. Successful exchange of messages does not guarantee that clinically useful transfer of information will result and the real objective of message exchange is to achieve clinically useful information transfer.

On the other hand, a test for interconnectivity between two Application Entities is not intended to be a detailed test process for debugging either or both of the Application Entities involved. Testing should be kept at a level that will demonstrate successful or unsuccessful inter-operation without going down to a level to ‘debug’ why the two Application Entities fail to inter-operate successfully.

The scope of the test plan will include functionality testing of the Application Entities, the operator interfaces for invoking message transfer, and indications of successful and failed operations.

2.4	References.  References to the following documents, when they exist, are required in the test plan:

Project Authorization

Project Plan

Quality Assurance Plan

Configuration Management Plan

Relevant Policies

Relevant Standards

In multilevel test plans, each lower-level plan must reference the next higher-level plan.

The following documents shall be included in the references of every test for interconnectivity of two Application Entities:

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statements of ecah of the Application Entities to be tested.

3.	Test Items

The items to be tested for interconnectivity will be the two Application Entities of interest. These test items shall be identified down to the level of version/revision number.

As specified by the DICOM Standard, the Upper Layer Association Services are included in the Application Entities and are therefore included in the testing process. In this test plan for the Storage Service Class, the assumption is that testing of the Upper Layer Association Services will have been completed prior to initiating testing for the Storage Service Class.

The following documents, if available, will be used to to provide the definition of correct operation.

Requirements Specification

Design Specification

User’s Guide

Operations Guide

Installation Guide

The following documents shall be included in the references of every test for operability of two Application Entities:

DICOM Standard, ACR/NEMA Standards Publication No. PS 3

The Conformance Statement of each Application Entity to be tested.

4.	Features to be Tested

The following features will be tested in every test of  Storage Service Class interconnectivity between two Application Entities. Storage Service will be tested for all IODs mutually supported by the two Application Entities under test.	The tests will conducted for all the roles that each of the Application Entities can assume. That is, if each Application Entity can be both a Service Class User and a Service Class Provider, then the tests should be conducted in both directions.

5.	Features Not to be Tested

The test of inter-operation of the Storage Service Class of two Application Entities is not a test of  the Association Services of the two Application Entities. The assumption is that testing of the Upper Layer Association Services and association negotiation will have been completed prior to initiating tests of the Storage Service Class. Furthermore, inter-operation testing for the Storage Service Class is not a test of the other service classes of the DICOM Standard, PS 3.

6.	Approach

The test personnel will use the DICOM Standard, PS 3, the respective Conformance Statements and documentation of the Application Entities to prepare all test design, test case, and test procedure specifications.

The overall test plan and the design of tests will depend on the the amount of information and the control that the designers and implementers have made available to the system operators. In testing for interconnectivity of the Storage Service Class, there are several levels of information access that may be available to the user:

Event Access:	Access to the information that a significant event has occurred.

Command Data Access:	Access to the relevant portions of the command set information.

Data Set Access:	Access to the information contained in the data set of the messages.

The assumption made in this template version of the test plan is that at least Event Access and Command Data Access levels of information are available for testing. The approach taken for each test will be to run the test and to ensure first that the significant primitive events have taken place in the proper sequence. Then the Command Data level can be examined and, finally,  the Data Access level of information, if available, can be examined.

7.	Item Pass/Fail Criteria

The Pass/Fail criterion for Storage Service Class interconnectivity of two Application Entities is that they should exhibit the behavior expressed in the DICOM Standard, PS 3, Part 4 and Part 7. For each invocation of the storage service, the primitive operations should take place in the proper sequence and each of the Application Entities should be left in the appropriate state relative to the association. The Commamd Set data, if available, can be checked to confirm that both Application Entities indicate the same outcome.

8.	Suspension Criteria and Resumption Requirements

8.1 Suspension Criteria.	Should the storage operation not perform as predicted for any of the Information Object Definitions that the Application Entities claim to support, then the tests of the Storage Service Class shall be suspended.

8.2 Resumption Requirements.	Testing of the Storage Service Class for the two Application Entities in question will resume from the beginning when the reason for suspension have been determined in either Application Entity, have been corrected, and new versions of either or both Application  Entities have been submitted to the testing group.

9.	Test Deliverables

The following documents should be included:

Test Plan

Test Design Specifications

Test Case Specifications

Test Procedure Specifications

Test Item Transmittal Reports

Test Logs

Test Incident Report Logs

Test Summary Reports

Test Input Data and Test Output Data

Test Tools

10.	Testing  Tasks

Identify the set of tasks necessary to prepare for and to perform the tests. Identify all intertask dependencies and any special skills required.

Tasks include:

(1)	Preparation of the test plan

(2)	Preparation of test design specification

(3)	Preparation of each test-case specification

(4)	Preparation of test-procedure specification

(5)	Prepare the hardware test environment

(6)	Prepare the software test environment

(7)	Perform the test procedures

(8)	Resolve test incident reports

(9)	Repeat tasks (6) - (9) until all test procedures are successful

(10)	Prepare the test summary report

(11)	Tramsmit all test documentation and test data to the appropriate groups

11.	Environmental Needs

This information specifies both the necessary and desired properties of the test environment. This specification should contain all the pertinent characteristics of the testing facilities including: the hardware, the communications and system software, the mode of usage (for example, stand-alone), and any other software or facilities needed to support the test.

12.	Responsibilities

Identify the groups responsible for managing, designing, preparing, executing, witnessing, checking, and resolving all aspects of the testing process. In addition, identify the groups responsible for providing the test items and the environmental needs.

13.	Staffing and Training

Specify test staffing needs by skill level. Identify training options for providing necessary skills.

14.	Schedule

Identify the significant milestones of the testing project, including any test milestones specified in the Software Project Schedule.

15.	Risks and Contingencies

Identify the high-risk assumptions of the test plan. Specify contingency plans for each assumption.

16.	Approvals

Specify the names and titles of all responsible persons who must approve the test plan. Signatures and dates of approval are desireable.